## Slide 1: PREVENTION OF ACCIDENTAL EXPOSURE IN RADIOTHERAPY
Part 6: Case histories of major accidents with abandoned radiotherapy sources
IAEA Training Course

## Slide 2: Overview / Objectives
Module 6.1: Source not under control (Brazil)
Module 6.2: Source not under control (Mexico)
Module 6.3: Source not under control (Turkey and Thailand)
Group exercise G7 : Public exposure
Objectives: To review and analyze case histories of major accidents involving abandoned radiotherapy sources

## Slide 3: Module 6.1: Source not under control (Brazil)
IAEA Training Course

## Slide 4: Background information

## Slide 5: Background information
Year 1985, a private radio-therapy institute, Instituto Goiano de Radioterapia (IGR), had two external beam treatment units:
  Cobalt-60 teletherapy unit
  Caesium-137 teletherapy unit

## Slide 6
The institute was situated in Goiânia, population ~1M, in central Brazil.
Both the 60Co and 137Cs units had gone through proper licensing procedures with CNEN (Regulatory Authority)
Rio de Janeiro
Background information

## Slide 7
CNEN is the Regulatory Authority, operating a licensing system relating to individuals (certifying training) and facilities (licensing operation of facilities).
Federal Ministry of Health had responsibilities for subsequent inspections of medical facilities. This was also devolved to State Health Secretaries.
Background information
Notes: Responsibilities were divided as outlined on this slide.

## Slide 8
The radiation head is adjustable vertically and can be rotated about two horizontal axes
Inside the radiation head is a rotating assembly with the sealed 137Cs source
Description of the Cs-unit
Notes: As part of the background, the characteristics of the Cs unit are outlined on the following slides.

## Slide 9
A source wheel forms a rotating shutter mechanism with the source
To produce a beam, the shutter is rotated electrically to align the source with an aperture
Description of the Cs-unit

## Slide 10
Cross-section of international standard capsule
Source material was inside two stainless steel capsules, inside a standard capsule
Description of the Cs-unit

## Slide 11
Source material: 137Cs chloride salt (which is highly soluble) – 93 g
Sep. 1987: 50.9 TBq; 4.56 Gy/h at 1 m
Description of the Cs-unit
Notes: It is important to note the soluble characteristics of the source material due to the role this plays in the event.

## Slide 12: Description of the event

## Slide 13
At the end of 1985, the IGR ceased operations at the old site and a new partnership took over the old site
The 60Co unit was moved to the new site
Start of the event

## Slide 14
Ownership of the contents of the old site became disputed
The 137Cs unit was left behind at the old site
CNEN did not receive appropriate notification of these changes in status
Start of the event

## Slide 15
Most of the old site was demolished.
The treatment rooms were not demolished but were left in a derelict state and apparently used by vagrants.
Start of the event

## Slide 16
NB! Original drawing made at time of discovery differs in details from description of event
Chronology of the event
Notes: When describing the chronology of the event in the following slides, the pictures are based on the original drawing in this slide, made just after the discovery of the event. It was later shown that the chronology in the original drawing differed somewhat to the actual chronology. The case history is gone through in great detail, in order to illustrate how an event of this type can unfold.

## Slide 17
10-13 Sep. 1987
RA and WP went to old site of IGR on rumours that valuable equipment had been left behind, and tried to dismantle Cs-unit with simple tools.
They succeeded in removing rotating assembly in its stainless steel casing.
Chronology of the event
Notes: Names are presented as initials in this case history, as in the reference publication.

## Slide 18
10-13 Sep. 1987
RA and WP took the rotating assembly in a wheelbarrow to RA’s house.
{no contamination was found at clinic – source assembly probably still intact}
{they would potentially have been exposed to the direct beam}
Chronology of the event
Notes: Within brackets {two last points in this slide} are some comments inserted, in order to not disturb the flow of the case history.

## Slide 19
13-15 Sep. 1987
WP and RA were vomiting, assuming this was due to something eaten
WP had diarrhoea, one hand swollen, sought medical assistance
Symptoms diagnosed as allergic reaction due to bad food
Chronology of the event

## Slide 20
13-18 Sep. 1987
Rotating assembly had been placed in RA’s yard, near houses rented out by RA’s mother
RA worked to remove the source wheel intermittently
Chronology of the event

## Slide 21
18 Sep. 1987
RA succeeded in removing the source wheel eventually
Chronology of the event

## Slide 22
18 Sep. 1987
In the course of removing the source wheel, the 1 mm window of the source capsule was punctured with a screwdriver and some of the source was scooped out
{residual contamination (2 Oct) under the mango tree gave a dose rate of 1.1 Gy/h at 1m}
Chronology of the event

## Slide 23
18 Sep. 1987
Rotating assembly pieces were sold to junkyard manager DF. Pieces were transported by employee of DF to garage in junkyard
That night, DF noticed a blue glow from the source capsule. He thought the powder might be valuable
Chronology of the event

## Slide 24
18-21 Sep. 1987
DF took the capsule into the house. Over the next three days, various neighbours, relatives and acquaintances were invited to see the capsule as a curiosity
Chronology of the event

## Slide 25
18-21 Sep. 1987
DF and his wife MF1 examined the powder closely
{MF1 (D=5.7 Gy) subsequently died. DF (D=7.0 Gy) survived, possibly due to fractionation arising from him being in and out of the house}
Chronology of the event

## Slide 26
21 Sep. 1987
EF1 visited DF and removed fragments of source from capsule. EF1 gave some fragments to his brother EF2 and took the rest home.
DF also distributed fragments to his family. Some persons applied powder on skin as glitter.
Chronology of the event

## Slide 27
21-23 Sep. 1987
MF1 was vomiting and had diarrhoea. After examination in hospital, MF1 was sent home and her mother MA1 came over for two days to nurse her
MA1 returned home on the bus, taking contamination with her
{MA1 (D=4.3 Gy) survived}
Chronology of the event

## Slide 28
22-24 Sep. 1987
DF’s employees, IS and AS worked on the rotating assembly with the unshielded source to extract lead. ZS visited, offered to cut up pieces with torch, but forgot to do so
{IS (D=4.5 Gy) and AS (D=5.3 Gy) subsequently died}
Chronology of the event

## Slide 29
24 Sep. 1987
IF, the brother of DF, was given source fragments, took them home, and placed them on the table during the meal
His six year old daughter, LF2, handled them while eating
{LF2 (D=6.0 Gy) subsequently died}
Chronology of the event

## Slide 30
24-28 Sep. 1987
The parts spread to two more junkyards
By 28 Sep. a significant number of people were physically ill
Chronology of the event

## Slide 31
28 Sep. 1987
MF1 was convinced that the glowing powder was causing the sickness
Source assembly was taken to small clinic in bag, left to Dr. PM, who got worried enough to put bag in yard
Dr. AM at Toxicological Information Centre was contacted
Chronology of the event

## Slide 32
28 Sep. 1987
Dr. JP at the State Department of the Environment was in turn contacted. He proposed that a Medical Physicist should have a look at the suspicious package
The pace of the events then quickened as the seriousness of the accident began to be appreciated
Chronology of the event
Notes: The event takes a leap here, where it is felt that a medical physicist should be contacted to investigate the suspicions.

## Slide 33
29 Sep. 1987
08:00-11:00
Medical Physicist WF was contacted and became convinced (from scintillation detector readings) that a major radiation source was at the clinic
The clinic was vacated. Police and fire brigade stopped building being entered
Chronology of the event
Notes: Starting with this slide, the events are described at different times on the same day. The chain of events becomes much more rapid. WF first came out to the clinic, turned on the meter, and found a full scale reading irrespective of the direction he was pointing the meter in. Probably considering that there might be a problem with the meter, he went back for another replacement meter. Returning to the clinic, he had the foresight to switch this meter on already when leaving the office, thus making it possible for him to rule out a meter-problem when the replacement meter again displayed a full scale reading having returned to the clinic.

## Slide 34
29 Sep. 1987
12:00-15:00
Radiation monitor showed contamination at first junkyard which was vacated
State Secretary for Health was informed of incident and its significance and to obtain further assistance
Director of the Department of Nuclear Installations in CNEN (coordinator for nuclear emergencies) was contacted
IGR was contacted, who tentatively identified source as possibly originating from IGR
Chronology of the event

## Slide 35
29 Sep. 1987
16:00-20:00
Hospitals were informed of potential radiation exposure to a number of people
Civil defence forces were alerted
Known sites of contamination were resurveyed
State Health Secretary made plans for receiving contaminated persons in city’s stadium
Press was taking an interest
More sites of major contamination were identified
Chronology of the event

## Slide 36
Chronology of the event
The physicist WF monitoring for contamination at the stadium

## Slide 37: Extent of the accident
Notes: Starting with this slide, the focus moves to study the magnitude of the event.

## Slide 38
112,000 persons were monitored
249 of these had external / internal contamination (up to 7 Gy)
129 of these had both external and internal contamination
49 of these were admitted to hospital
20 of these needed intensive medical care
10 of these were in critical conditions
4 of these died (within four weeks) and one had to have the forearm amputated
Extent of the accident
Notes: This shows the extent of the accident in terms of persons affected by contamination.

## Slide 39
Extent of the accident
Notes: The persons most highly contaminated in the accident in Goiania, listed by site of exposure and family membership. Estimates from cytogenetic data of doses incurred (in grays) and information on those admitted to hospital and on the four fatalities are also given. (A thick black border indicates who was hospitalised; the other individuals were treated as out-patients.)

## Slide 40
7 main foci of contamination within 1 km2
Dose rates up to 2 Sv/h at 1 m from contamination
41 houses were evacuated
Contamination was removed from 45 different public places, and it was also found on 50 vehicles
Extensive rains dispersed the highly soluble caesium chloride into the environment
Extent of the accident
Notes: This shows the extent of the accident in terms of environment affected by contamination.

## Slide 41: Actions taken
Notes: Starting with this slide, the actions taken in response to the accident are presented.

## Slide 42
Remember the cardinal rule of radiation protection: The security of the source is of paramount importance!
This was an Emergency Exposure Situation (BSS). The Regulatory Authority CNEN was contacted
Civil defence (police, etc.) was contacted
Intervention was enacted based on action levels (e.g. evacuation, decontamination, removal of soil)
Actions taken
Notes: It is possible to refer back to the terminology contained in the International BSS at this stage.

## Slide 43
For accidents like this, the actions to be taken can be divided into two phases:
The initial phase, when urgent action is required
  To identify potential sources of acute exposure
  To bring exposure under control
The recovery phase, when urgent action is no longer required and the objective is to restore the situation to normal
Actions taken

## Slide 44
Initial phase
Identification of main contamination sites
Evacuation at sites above intervention levels
Control areas through access prevention
Identification of persons who had incurred significant doses or were contaminated
+ Medical response, of course!
Actions taken

## Slide 45
Medical response
Specialists dispatched to Goiânia
Severe external and internal contamination with Cs-137
External decontamination was performed
“Prussian Blue” was used for internal decontamination
Acute radiation syndrome and local injuries were treated
Actions taken
Notes: It is important to point out how the soluble characteristics of the source in this case made internal contamination a feature of the event.

## Slide 46
Remedial actions
Decontamination of property
Collection of contaminated clothing
Removal of contaminated soil
Placing of restrictions on home grown produce
Actions taken

## Slide 47
Actions taken
Seven houses had to be demolished
Preparing to demolish the house of EF2 near first junkyard

## Slide 48
Removal of waste material
The final total volume of waste stored was 3500 m3
This represents more than 275 lorry loads
Since action levels were chosen restrictively, in relation to international guidelines, this volume became very large and the operation very costly
Actions taken

## Slide 49: Lessons and recommendations

## Slide 50: Initiating event and contributory factors
The event was triggered by
  Allowing a source to become in a state of not being secure and under control
Contributory factors
  The solubility and ease of dispersion of the caesium chloride

## Slide 51
The person designated as being liable for a radioactive source must always ensure the source is secure and under control
Notify Regulatory Authorities of changes in circumstances relating to sources
Physical and chemical properties of sources should be taken into account in licensing
Lessons for prevention
Notes: Compare to BSS.

## Slide 52
Too restrictive criteria for action levels (in relation to international guidelines) can impose a tremendous burden on public and response team
Recognition of the nature of radiation injury depends on education of health and safety professions. This can aid speed of diagnosis and overall speed of response
Prompt handling of internal contamination is important. Prussian Blue was shown to aid under these circumstances
Lessons for responses

## Slide 53
Caesium-137 source not under control
Many people lost their homes
4 persons dead
Most fatal radiological accidents have occurred outside the nuclear industry
Goiâna accident summary

## Slide 54: Reference
IAEA: The Radiological Accident in Goiânia (1988)
rpop.iaea.org