## Slide 1: Module 2.10: Accident update – some newer events (UK, USA, France)
IAEA Training Course

## Slide 2
Questions
Do you think the accidents have not happened in recent years?
Do you think well-developed centres are immune to these accidents?
Notes: The answers will come at the end of the lecture…

## Slide 3
Overview
It should be noted that the intent is certainly not to reflect the quoted centres in this presentation in poor light
Instead, the purpose is to draw lessons
In many cases, the centres have a quality system in place
The events are reconstructed from information in the public domain, and might differ from actual events due to gaps in this information.

## Slide 4
Overview
1st example: Incorrect manual parameter transfer (UK)
2nd example: Reversal of images (USA)
3rd example: Inappropriate measuring device (France)
4th example: Erroneous calculation for soft wedges (France)
5th example: Incorrect IMRT planning (USA)
6th example: More information needed…
Newer examples of accidents in radiotherapy from 2004 to 2007

## Slide 5: 1st example: Incorrect manual parameter transfer (UK)
IAEA Training Course

## Slide 6: Background
January 2006 at the Beatson Oncology Centre (BOC) in Glasgow, Scotland
  At the time: Radiotherapy physics staffing levels in Scotland less than 60% of the recommended level
  “Glasgow has problems with recruiting physicists, as shown by their high number of vacancies.”
The Beatson Oncology
Centre in Glasgow
Notes: <footer>
It will be argued later that staffing pressures were identified as a significant underlying cause of the incident. The statement that “Glasgow has problems with recruiting physicists…” is taken from a review: “‘Equipment, Workload and Staffing for Radiotherapy in Scotland 1997–2003” (Board of the Faculty of Clinical Oncology, Royal College of Radiologists, London), 2005.

## Slide 7
Background
Treatment planning at BOC:
  14.5 whole time equivalent (WTE) staff were available for between 4500 and 5000 new treatment plans per year.
  When staffing levels were compared with guidelines from IPEM, it was seen that 18 WTE staff would be the recommended level.
Notes: IPEM is the UK Institute of Physics and Engineering in Medicine; ”Guidelines: The provision of a physics service to radiotherapy” (2002).

## Slide 8
Background
Treatment planning at BOC:
  Planning staff members and planning procedures were both categorized
  A to C denotes senior to junior staff
  A to E denotes simple to complex plans
  The main duties per staff category is outlined in column 4
Table from: “Report of an investigation by the Inspector appointed by the Scottish Ministers for The Ionising Radiation (Medical Exposures) Regulations 2000”
Notes: It is later shown that a complex plan was planned by a more junior category of staff than intended.

## Slide 9
Background
Treatment planning at BOC:
  Practice prior to 2005 had been to let the treatment planning system (TPS) calculate the Monitor Units (MU) for 1 Gy followed by manual multiplication with the intended dose per fraction for the correct MU-setting to use.
Notes: The reasoning behind inputting dose per fraction already in the TPS, is based on avoiding manual errors in the multiplication with the intended dose per fraction in order to get the correct number of monitor units to use.

## Slide 10
Background
Treatment planning at BOC:
  In May 2005, the Record and Verify (RV) system was upgraded to be a more integrated platform.
  The centre decided to input the dose per fraction already in the TPS, for most but not all treatment techniques.
Notes: The reasoning behind not introducing this for all treatment techniques is not clear from the reports, but could, in the case of medulloblastoma treatment, be due to the prescribed dose varying during treatment and due to spine field compensators being calculated without the TPS.

## Slide 11: What happened?
5th January 2006, Lisa Norris, 15 years old, started her whole CNS treatment at BOC
The treatment plan was divided into head-fields and lower and upper spine-fields
This is considered to be a complex treatment plan, performed about six times per year at the BOC.
Lisa Norris
Notes: There are several pictures, as well as her name, of Lisa Norris in the public domain. Therefore her picture and name also appears in this lecture.

## Slide 12
What happened?
The bulk of the planning was done by “Planner X” in Dec’05, a junior planner
“Planner X” had not yet been registered internally to be competent to plan whole CNS, or to train on these
“Planner X” got initial instructions and the opportunity to be supervised when creating the plan
Notes: ”Planner X” did not ask for supervision. He was unaware of the problems that were present in the planning approach he was taking. The supervision seems to have been reactive in nature. CNS = central nervous system

## Slide 13
What happened?
Whole CNS plans still went by the “old system”, where TPS calculates MU for 1 Gy with subsequent upscaling for dose per fx
A “medulla planning form” was used, which is passed to treatment radiographers for final MU calculations
Table from: “Report of an investigation by the Inspector appointed by the Scottish Ministers for The Ionising Radiation (Medical Exposures) Regulations 2000”
Notes: TPS = treatment planning system

## Slide 14
What happened?
HOWEVER – “Planner X” let the TPS calculate the MU for the full dose per fx – not for 1 Gy as intended
Since the dose per fx to the head was 1.67 Gy, the MU’s entered in the form were 67% too high for each of the head-fields
Table from: “Report of an investigation by the Inspector appointed by the Scottish Ministers for The Ionising Radiation (Medical Exposures) Regulations 2000”
Notes: Fx = fraction; NB: The dose per fraction to the head was really supposed to be 175 cGy, but there also seems to have been a mistake in the total number of fractions to give. If this is emphasized too much in the lecture, it might take the focus away too much from the main lessons to learn.

## Slide 15
What happened?
This error was not found by the more senior planners who checked the plan
The radiographer on the unit thus multiplied with the dose per fx a second time
2.92 Gy per fx to the head
Table from: “Report of an investigation by the Inspector appointed by the Scottish Ministers for The Ionising Radiation (Medical Exposures) Regulations 2000”
Notes: 2.92 Gy comes from 1.67 x 1.75 Gy (see notes on previous slide)

## Slide 16: Discovery of accident
“Planner X” calculated another plan of the same kind and made the same mistake
This time, the error was discovered by a senior checker (1st of Feb ‘’06)
The same day, the error in calculations for Lisa Norris was also identified

## Slide 17
The total dose to Lisa Norris from the Right and Left Lateral head fields was 55.5 Gy (19 x 2.92 Gy)
She died nine months after the accident
Impact of accident

## Slide 18: Lessons to learn
Ensure that all staff
  Are properly trained in safety critical procedures
  Are included in training programmes and has supervision as necessary, and that records of training are kept up-to-date
  Understand their responsibilities
Include in the Quality Assurance Program
  Formal procedures for verifying the risks following the introduction of new technologies and procedures
  Independent MU checking of ALL treatment plans
Review staffing levels and competencies

## Slide 19: References
Unintended overexposure of patient Lisa Norris during radiotherapy treatment at the Beatson Oncology Centre, Glasgow in January 2006. Report of an investigation by the Inspector appointed by the Scottish Ministers for The Ionising Radiation (Medical Exposures) Regulations 2000 (2006)
Cancer in Scotland: Radiotherapy Activity Planning for Scotland 2011 – 2015. Report of The Radiotherapy Activity Planning Steering Group’ The Scottish Executive. Edinburgh. (2006)
The Glasgow incident – a physicist’s reflections. W.P.M. Mayles. Clin Oncol 19:4-7 (2007)
Radiotherapy near misses, incidents and errors: radiotherapy incident in Glasgow. M.V. Williams. Clin Oncol 19:1-3 (2007)

## Slide 20: 2nd example: Reversal of images (USA)
IAEA Training Course

## Slide 21
What happened?
October 2007 at the Karmanos Cancer Center (KCC) in midtown Detroit, Michigan, USA
  At the Gamma Knife treatment facility, a patient was set up for MRI imaging
  Standard practice is to position the patient “head first”
  The patient was positioned “head first”, but “feet first” scan technique was chosen on the unit
The KCC in Detroit

## Slide 22
What happened?
The axial images were therefore reversed left-to-right
  The physicist did not see the mistake when importing images into the TPS
  The error resulted in an 18 mm shift of isocentre across the midline of the brain
Stereotactic treatment
(image from KCC)

## Slide 23: Lessons to learn
Include in the Quality Assurance Program
  Procedures for verifying left from right in safety critical images, e.g. by using fiducial markers
  Ensure there are written protocols posted, known and followed, for safety critical procedures

## Slide 24: References
Gamma knife treatment to wrong side of brain. Event Notification Report 43746. United States Nuclear Regulatory Commission (2007)

## Slide 25: 3rd example: Inappropriate measuring device (France)
IAEA Training Course

## Slide 26
Background
Reported 2007 at Hôpital de Rangueil in Toulouse, France
  In April 2006, the physicist in the clinic commissioned the new BrainLAB Novalis stereotactic unit
  This unit can operate with microMLC’s (3 mm leaf-width) or conical standard collimators
The Hôpital de Rangueil
in Toulouse
Notes: microMLC = micro multileaf collimators, with very thin leaf width, capable of shaping small fields with high resoultion

## Slide 27
Background
Very small fields can be defined with the microMLC’s
  High dose to a 6 x 6 mm field is within capability
  The TPS requires percent depth doses, beam profiles and relative scatter factors down to this field size
  Care must be taken when measuring small fields!

## Slide 28
What happened?
Different measuring devices were used by the physicist
  A measuring device not suitable for calibrating the smallest microbeams was used
  “…an ionisation chamber of inappropriate dimensions…” according to Nuclear Safety Authority (ASN) inspectors

## Slide 29
What happened?
The incorrect data was entered into the TPS
  All patients treated with micro MLC were planned based on this incorrect data
  Patients treated with conical collimator were not affected

## Slide 30
BrainLAB discovered that the measurement files did not match up with those at other comparable centres, during a worldwide intercomparison study
It should be noted that the company does not validate or hold responsibility for local measurements or implementation
Discovery of accident

## Slide 31
Impact of accident
Treatment based on the incorrect data went on for a year (Apr´06 – Apr´07)
All patients treated with microMLC were affected (145 of 172 stereotactic patients)
The dosimetric impact was evaluated as small in most cases, with 6 patients identified for whom over 5% of the volume of healthy organs may have been affected by dose exceeding limits

## Slide 32: Lessons to learn
Ensure that staff
  Understand the properties and limitations of the equipment they are using
Include in the Quality Assurance Program
  Intercomparison with other hospitals, i.e. independent check of new equipment by independent group (using independent equipment) before equipment is clinically used

## Slide 33: References
Report concerning the radiotherapy incident at the university hospital centre (CHU) in Toulouse – Rangueil Hospital. ASN – Autorité de Sûreté Nucléaire (2007)

## Slide 34: 4th example: Erroneous calculation for soft wedges (France)
IAEA Training Course

## Slide 35
Background
In May 2004 at Centre Hospitalier Jean Monnet in Epinal, France
  …it was decided to change from static (hard) wedges to dynamic (soft) wedges for prostate cancer patients
  In a country of few Medical Physicists (MP), this facility had a single MP who was also on call in another clinic
The Jean Monnet Hospital
in Epinal
Notes: microMLC = micro multileaf collimators, with very thin leaf width, capable of shaping small fields with high resoultion

## Slide 36
Background
In preparation for the change in treatment technique, two operators (treatment planners?) were given two brief demo’s
  The operators did not have any operating manual in their native language

## Slide 37
Background
When the soft wedges were introduced:
  The independent MU check in use could not be used anymore (unless modified)
  The diodes used for independent dose check could not be correctly interpreted anymore
Notes: It is possible to modify an in-house constructed MU checker to take into account dynamic wedges. A commercial program should already have this option included. Likewise, it is possible to estimate predicted diode readings for treatment with dynamic wedges.

## Slide 38
What happened?
Treatment planning with soft wedges started
  Not all the treatment planners did understand the interface to the planning system

## Slide 39
What happened?
Treatment planning with soft wedges started
  Not all the treatment planners did understand the interface to the planning system
  Some selected the planning for mechanical wedge when intending dynamic wedge
v

## Slide 40
What happened?
Treatment planning with soft wedges started
  Not all the treatment planners did understand the interface to the planning system
  Some selected the planning for mechanical wedge when intending dynamic wedge
  Instead they should have selected Dynamic Wedge…
v

## Slide 41
What happened?
Treatment planning with soft wedges started
  Not all the treatment planners did understand the interface to the planning system
  Some selected the planning for mechanical wedge when intending dynamic wedge
  Instead they should have selected Dynamic Wedge…
    …which would have let the correct planning tool appear
v

## Slide 42
What happened?
When planning was finished and the isodose distribution approved
  …the parameters were manually transferred to the treatment unit
  Manually transferred MU’s would have been calculated for mechanical wedges and would be much greater than what is needed for giving the same dose with dynamic wedges

## Slide 43
Details not clear, BUT: it might have been when MU check software was replaced and updated to be able to handle independent checking of dynamic wedges.
Discovery of accident
Notes: It is difficult to find the exact details on how the discovery was made. In the official report it says: ”During the summer of 2005, one of the radiotherapists and the medical physicist discovered the overdosage error by re-examining all the files”. It also states that ”From May 2005, five patients were suffering from severe lesions; the number rose to seven in June 2005 , and to ten by August 2005”. The report doesn’t explicitly, however, make the link between the unwell patients and the re-examination of all files. In another part, the report states: ”The overdosage…took place…to 1 August 2005. After this date, new dosimetry software replaced the old software, whose ergonomics was incapable of preventing this type of error.”

## Slide 44
Impact of accident
Treatment based on incorrect MU’s went on for over a year (6 May 2004 – 1 Aug 2005)
At least 23 patients received overdose (20% or more than intended dose)
Between September 2005 and September 2006, four patients died. At least ten patients show severe radiation complications (symptoms such as intense pain, discharges and fistulas)

## Slide 45
Information following accident
15 Sep 2005, two doctors from the clinic passed on information that went to the Regional Dept. of Health and Social Security (DDASS)
5 Oct 2005 a meeting was held at DDASS. Decisions were not documented or uniformly interpreted.
National authorities in charge were not informed at this stage, but only a full year after the accident (July 2006)

## Slide 46
Information following accident
7 patients were informed during the last quarter of 2005.
16 other patients were (wrongly) considered no to be affected. Of these …
    … 3 were informed by another doctor than their radiotherapist
    … 1 learnt from a third party person
    … 1 learnt from the press
    … 1 learnt by overhearing a doctor speaking to a colleague
    … 4 were informed by management 2 days before press release
    … 1 died before being informed
Notes: This information is from the official report by ASN.

## Slide 47: Lessons to learn
Ensure that staff
  Understand the properties and limitations of the equipment they are using
  Are properly trained in safety critical procedures
Include in the Quality Assurance Program
  Formal procedures for verifying new technologies and procedures before implementation
  Independent MU checking of ALL treatment plans
  In vivo dosimetry
Make sure the clinic has a system in place for
  Investigation and reporting of accidents
  Patient management and follow up, including communication to patients
Instructions should be in a language that is understood

## Slide 48
References
Summary of ASN report n° 2006 ENSTR 019 - IGAS n° RM 2007-015P on the Epinal radiotherapy accident. G. Wack, F. Lalande, M.D. Seligman (2007)
Accident de radiothérapie à Épinal. P.J. Compte. Société Française de Physique Médicale (2006)
Lessons from Epinal. D. Ash. Clin Oncol 19:614-615 (2007)

## Slide 49
Postscript to accident in Epinal
Going through the records, two further episodes were reported subsequently
Reported in Feb 2007:
    In the time period 2001-2006, portal imaging was used repeatedly without taking into account the added dose (estimated to have been +8% of total) for 412 patients under medical survey
Reported in July 2007:
    In the time period 1989-2000, use of an in-house TPS not updated after change in treatment technique, might have led to 300 patients receiving up to 7% added dose.

## Slide 50: 5th example: Incorrect IMRT planning (USA)
IAEA Training Course

## Slide 51
Background
March 2005, somewhere in the state of New York, USA
  A patient is due to be treated with IMRT for head and neck cancer (oropharynx)
Notes: The oropharynx-pointer is underlined in red. IMRT = intensity modulated radiation therapy

## Slide 52
What happened?
March 4 – 7, 2005
  An IMRT plan is prepared: “1 Oropharyn”. A verification plan is created in the TPS and measurements by Portal Dosimetry (with EPID) confirms correctness.
Example of an EPID (Electronic Portal Imaging Device) (Picture: P.Munro)
Notes: A verification plan is used to verify that the calculated dose distribution is achieved at irradiation. In this case, the EPID is irradiated and used as a dosimeter. The measured dose distribution is subsequently compared with the calculated dose distribution.

## Slide 53
What happened?
March 8, 2005
  The patient begins treatment with the plan “1 Oropharyn”. This treatment is delivered correctly.
“Model view” of treatment plan (Picture: VMS)
Notes: The picture of the ”model view” shows the field outlines from various gantry angles.

## Slide 54
What happened?
March 9-11, 2005
  Fractions #2, 3 and 4 are also delivered correctly. Verification images for the kV imaging system are created and added to the plan, now called “1A Oropharyn”.
“Model view” of treatment plan (Picture: VMS)
Notes: Verification images are digitally reconstructed radiographs showing the correct set-up location, for comparison with the images from the kV imaging system attached to the gantry of the linear accelerator.

## Slide 55
What happened?
March 11, 2005
  The physician reviews the case and wants a modified dose distribution (reducing dose to teeth) “1A Oropharyn” is copied and saved to the DB as “1B Oropharyn”.
“Model view” of treatment plan (Picture: VMS)
Notes: DB = database

## Slide 56
What happened?
March 14, 2005
  Re-optimization work on “1B Oropharyn” starts on workstation 2 (WS2).
  Fractionation is changed. Existing fluences are deleted and re-optimized. New optimal fluences are saved to DB.
  Final calculations are started, where MLC motion control points for IMRT are generated. Normal completion.
Multi Leaf Collimator (MLC)

## Slide 57
What happened?
March 14, 2005, 11 a.m.
  “Save all” is started. All new and modified data should be saved to the DB.
  In this process, data is sent to a holding area on the server, and not saved permanently until ALL data elements have been received.
  In this case, data to be saved included: (1) actual fluence data, (2) a DRR and (3) the MLC control points
A Digitally Reconstructed Radiograph (DRR) of the patient
Notes: The concept of the ”holding area” is important for what follows in the sequence of events.

## Slide 58
What happened?
March 14, 2005, 11 a.m.
The actual fluence data is saved normally.
  Next in line is the DRR. The “Save all” process continues with this, but is not completed.
  Saving of MLC control point data would be after the DRR, but will not start because of the above.
A Digitally Reconstructed Radiograph (DRR) of the patient

## Slide 59
What happened?
March 14, 2005, 11 a.m.
  An error message is displayed.
  The user presses “Yes”, which begins a second, separate, save transaction.
  MLC control point data is moved to the holding area.
The transaction error message displayed
Notes: The message on the ”transaction error message” can be read out. It can be said to be a typical example of this type of error messages, which are difficult to understand as an end-user, and where consequences from pressing either ”Yes” or ”No” are somewhat unclear.

## Slide 60
What happened?
March 14, 2005, 11.a.m.
  The DRR is, however, still locked into the faulty first attempt to save.
  This means the second save won’t be able to complete.
  The software would have appeared to be frozen.
The frozen state of the second “Save All” progress indication

## Slide 61
What happened?
March 14, 2005, 11.a.m.
  The user then terminated the TPS software manually, probably with Ctrl-Alt-Del or Windows Task Manager
  At manual termination, the DB performs a “roll-back” to return the data in the holding area to its last known valid state
  The treatment plan now contains (1) actual fluence data; (2) not the full DRR; (3) no MLC control point data
Ctrl-Alt-Del

## Slide 62
What happened?
March 14, 2005, 11.a.m.
  Within 12 s, another workstation, WS1, is used to open the patients plan. The planner would have seen this:
        Valid fluences were already saved. Calculation of dose distribution is now done by the planner and saved. MLC control point data is not required for calculation of dose distribution.
Sagittal view of patient, with fields and
dose distribution
Notes: It can be seen that MLC data are not included here since the sagittal views of the fields (yellow lines in the picture) are straight and indicate open fields. This might be difficult to see, though.

## Slide 63
What happened?
March 14, 2005, 11.a.m.
  No control point data is included in the plan.
The sagittal view should have looked like the one to the right, with MLCs
Notes: On the right-hand picture, the steps of the MLCs are clearly seen.

## Slide 64
What happened?
March 14, 2005, 11 a.m.
  No verification plan is generated or used for checking purposes, prior to treatment (should be done according to clinics QA programme)
  The plan is subsequently prepared for treatment (treatment scheduling, image scheduling, etc) – after several computer crashes.
  It is also approved by a physician
  According to QA programme, a second physicist should then have reviewed the plan, including an overview of the irradiated area outline, and the MLC shape used.
Notes: It is important to emphasize that the QA programme in the clinic would have had a good opportunity to prevent this from happening, had it been applied as planned.

## Slide 65
What happened?
Would have been seen on verification:
Notes: It is indicated in the text that MLC used is ”NONE” and in the DRR an open field is shown (yellow line). The reference does not indicate if a visual verification of the treatment was performed.

## Slide 66
What happened?
Should have been seen on verification:
Notes: It is indicated in the text that MLC used is ”DoseDynamic” and in the DRR a field with an MLC outline is shown (in blue).

## Slide 67
What happened?
March 14, 2005, 1 p.m.
  The patient is treated. The console screen would have indicated that MLC is not being used during treatment:
Notes: In the text, MLC is empty. In the image, an open field defined by collimators is shown. It can also be seen that the patient is treated two hours after planning, which can indicate time-pressure.

## Slide 68
What happened?
March 14, 2005, 1 p.m.
  Expected display:
Notes: In the text, MLC is ”Dynamic”. In the image, a field defined by MLCs is shown..

## Slide 69
Discovery of accident
March 15-16, 2005
  The patient is treated without MLCs for three fractions
  On March 16, a verification plan is created and run on the treatment machine. The operator notices the absence of MLCs.
  A second verification plan is created and run with the same result.
  The patient plan is loaded and run, with the same result.
Impact of accident
The patient received 13 Gy per fraction for three fractions, i.e. 39 Gy in 3 fractions

## Slide 70
  Do what you should be doing according to your QA program – the error could have been found through verification plan (normal QA procedure at the facility) or independent review
  Be alert when computer crashes or freezes, when the data worked on is safety critical
  Work with awareness at treatment unit, and keep an eye out for unexpected behaviour of machine
Lessons to learn

## Slide 71
[Treatment Facility] Incident Evaluation Summary, CP-2005-049 VMS. 1-12 (2005)
ORH Information Notice 2005-01. Office of Radiological Health, NYC Department of Health and Mental Hygien (2005)
References

## Slide 72: 6th example: More information needed
IAEA Training Course

## Slide 73
Background
Unfortunately, information that professionals could learn from is not readily available sometimes -
  There is the case in Florida (2004), where it seems that a stereotactic unit was miscalibrated, resulting in 50% higher dose than intended for 77 patients with brain tumours.
  We could have learnt a lot here…

## Slide 74: Lessons to learn
?
References
-
Notes: The point here is that since there are no references readily available, there are no lessons that we can learn from this accident.

## Slide 75
Questions
Do you think the accidents have not happened in recent years? ANSWER: NO! If YES, then think again!
Do you think well-developed centres are immune to these accidents? ANSWER: NO! If YES, then think again!