## Slide 1: PREVENTION OF ACCIDENTAL EXPOSURE IN RADIOTHERAPY
Part 3: Analysis of causes and contributing factors
IAEA Training Course

## Slide 2: Overview / Objectives
Module 3.1: External beam therapy
Module 3.2: Brachytherapy
Group exercise G4: Dissemination of course material
Objectives:
To analyze causes and contributing factors from available information on a collection of incidents and accidental exposures occurred

## Slide 3: Module 3.1: Other cases (external beam therapy)
IAEA Training Course

## Slide 4
Case histories
Previous lectures have given details on some major accidental exposures.
This lecture presents other accidental exposures, collected in ICRP Publication 86 and IAEA Safety Report Series No.17.
References
rpop.iaea.org

## Slide 5
Case histories
There are lessons to learn for most steps in the radiotherapy process.
A collection of accidental exposures
Notes: This table, copied from the ICRP publication on the subject, is listing the number of cases in the said publication for each of the different steps in the radiotherapy process (both external beam therapy and brachytherapy).

## Slide 6
Classification of accidental exposures
Cases are grouped by the steps in relation to the radiotherapy process.
Selected accidental exposures presented
13 case histories and specific lessons learned for external beam radiotherapy
6 case histories and specific lessons learned for brachytherapy
Generic lessons learned
Are there recurring themes in the lessons learned?
Overview of lecture
Notes: The lecture has two modules – 3.1 (above the line) and 3.2 (below the line). The intention of this part of the workshop is to start looking more systematically for “patterns” in the lessons learned.

## Slide 7
Selected accidental exposures will be presented, with a short case history and a brief summary of initiating event, consequences and lessons learned.
How is the “initiating event” presented?
As some act that occurred at some point in time, which eventually led to the accidental exposure of patients.
How are the “consequences” presented?
As a brief summary of how patients were affected by the accidental exposure (in terms of dose-deviation from intended dose).
How are the “lessons learned” presented?
As specific actions or layers of safety that could have stopped the initiating event from becoming an accidental exposure with consequences for the patients, had these actions been performed.
Overview of lecture

## Slide 8
Equipment problems
Maintenance of radiotherapy equipment
Calibration of external beams
Treatment planning and dose calculations
Simulation
Treatment set-up and delivery
There will be examples of accidental exposures from most classes
Classification of accidental exposures
Notes: There will be examples of accidents from all different steps in relation to the external beam therapy process. We will start looking for recurring themes in the lessons learned from these accidents.

## Slide 9
External beam radiotherapy
Cases 1 - 13
Notes: This first module contains the external beam examples.

## Slide 10: Equipment problems
Case 1
Notes: First step represents equipment problems.

## Slide 11
1. Loose wedge mounting mechanism
Wedge factors were measured during commissioning of a Cobalt unit, using the beam in the vertical position.
When the gantry was rotated 90° for treatments with a horizontal beam, a loose wedge mounting mechanism allowed wedge filters to shift.
As a result, the dose distribution and the central axis wedge factor were incorrect, with the dose to the patient too high across the beam for one horizontal machine position and too low for the other horizontal machine position.
Equipment problems

## Slide 12
Equipment problems
1. Loose wedge mounting mechanism
Initiating event:
Mechanical deficiency related to the wedge holder
Consequences:
Patients had deviations in dose of up to 8% from prescribed doses.
Lessons learned:
When commissioning a treatment unit, remember to make some measurements at other machine positions than vertical.
Notes: It can be seen here in the first case, that the lessons learned are presented in a very specific way, derived directly from the short case history in the preceding slide. Towards the end of the next module, these specific lessons learned will be generalized.

## Slide 13
Also remember lectures on:
Canada and USA, 1985-1987
(Six accidental exposures involving software problems
in several accelerators of the same type)
Poland, 2001
(Accelerator malfunction)
Equipment problems
Notes: The lecturer can recap lectures from the previous part of the workshop on major case histories, and especially the essence of lessons learned from these cases.

## Slide 14
Maintenance of
radiotherapy equipment

## Slide 15
Maintenance of radiotherapy equipment
Remember lecture on:
Spain, 1990
(Incorrect repair followed by insufficient communication)
Notes: This major case history is the only example from this step in the ICRP publication.

## Slide 16
Calibration of external beams
Cases 2 - 5

## Slide 17
Calibration of external beams
2. Incorrect use of a plane parallel chamber
A new physicist at a hospital used a pancake chamber to calibrate several electron beams.
A label on the chamber, placed by the previous physicist, indicated the side on which the beam should be incident. Although the previous physicist had used the chamber correctly, his labelling was incorrect and the new physicist used the chamber upside down in the beam.
This resulted in wrong calibration, progressively worse for lower electron energies. The discrepancies were eventually revealed through mailed TLD-dosimetry.

## Slide 18
Calibration of external beams
2. Incorrect use of a plane parallel chamber
Initiating event:
Incorrect use of chamber for calibration
Consequences:
Some patients received the wrong dose per fraction (up to 20% overdose) before the error was corrected.
Lessons learned:
Make sure that instruments used are well understood in terms of how they work.
The physicist should take responsibility for all aspects of dosimetry.

## Slide 19
Calibration of external beams
3. Error in correction for atmospheric pressure
Atmospheric pressure is used for correcting dose measurements when using some dosimeters.
Four institutions were using atmospheric pressure data from nearby weather stations.
The physicists concerned did not realize that these data were actually corrected to sea-level, and thereby did not reflect the true value of atmospheric pressure at the institutions.
As a result, pressure correction-factors were incorrect, leading to incorrect calibrations.

## Slide 20
Calibration of external beams
3. Error in correction for atmospheric pressure
Initiating event:
Incorrect pressure values were used for measurement corrections.
Consequences:
Patients at these institutions received overdose of between 13% and 21%, in one case for ten months.
Lessons learned:
Have a functioning barometer in the institution, and know how to use it.
If requesting pressure-values from other source, make sure it is known what the data is referring to.

## Slide 21
Calibration of external beams
4. Dosimeter calibration report used incorrectly
An institution had its ionization chamber and electrometer calibrated for Cobalt-60 at a standards dosimetry laboratory.
The calibration certificate was in terms of dose to water, but was interpreted by the physicist at the institution as specifying dose in air.

## Slide 22
Calibration of external beams
4. Dosimeter calibration report used incorrectly
Initiating event:
The calibration certificate was used incorrectly.
Consequences:
Patients received ~ 11% overdose for at least one year.
Lessons learned:
Make sure you understand the calibration certificate.
Have another physicist calibrate the beam independently.

## Slide 23
Calibration of external beams
5. Incorrect calibration of a machine with asymmetric jaws
A linear accelerator with asymmetric jaws was calibrated with the detector positioned in the penumbra region.
The measured value at this position did not represent the dose in the centre of the field.

## Slide 24
Calibration of external beams
5. Incorrect calibration of a machine with asymmetric jaws
Initiating event:
The calibration of the beam was made in the penumbra.
Consequences:
Patients received an overdose of 27%.
Lessons learned:
Make sure you understand the features of an asymmetric beam.
Have another physicist calibrate the beam independently.

## Slide 25
Calibration of external beams
Also remember lectures on:
USA, 1974-76
(Incorrect 60Co decay chart and lack of verification)
Costa Rica, 1996
(Beam miscalibration following the exchange of a 60Co source)
France, 2006-2007
(Inappropriate measuring device)

## Slide 26
Treatment planning and dose calculation
Cases 6 - 10

## Slide 27
Treatment planning and dose calculation
6. Incorrect basic data in a TPS
Basic data used in a TPS was entered into the computer by a physicist. The input data differed from measured data for a particular linear accelerator.
The inconsistency was not detected during commissioning of the planning system.
A new physicist was appointed when the old physicist left. The reason for the errors remained unknown.

## Slide 28
Treatment planning and dose calculation
6. Incorrect basic data in a TPS
Initiating event:
Incorrect basic data entered into TPS.
Consequences:
Patients received a 15% overdose.
Lessons learned:
Make sure TPS is commissioned fully.
Check treatment plans independently.
Also measure basic treatment unit data and check against treatment planning data occasionally.

## Slide 29
Treatment planning and dose calculation
7. Incorrect depth dose data
A manufacturer was contracted to measure depth dose data during installation of a linear accelerator.
The local physicist later checked the data and found an 8% discrepancy for some field sizes and depths. He concluded that the manufacturer’s data were correct and used them clinically.
An outside consultant physicist later found that the measurements of the local physicist were correct. This was several months later.

## Slide 30
Treatment planning and dose calculation
7. Incorrect depth dose data
Initiating event:
Incorrect basic data tables for dose calculations created.
Consequences:
Some patients (over several months) received an 8% lower dose than prescribed.
Lessons learned:
Commission tables thoroughly before accepting to use them for treatment.
Resolve why there are discrepancies in data.
The physicist should take responsibility for all aspects of dosimetry.
Notes: It is important to emphasize to the participants that it is always the local physicist who is responsible for dosimetry – even when measurements have been performed by other parties.

## Slide 31
Treatment planning and dose calculation
8. Inconsistent sets of basic data
An institution had two sets of basic data available for clinical use, for one particular treatment unit (output factors, %dd, etc).
The two sets of data differed by 10%, with one set being correct.
These sets of data were used interchangeably for a period of time.

## Slide 32
Treatment planning and dose calculation
8. Inconsistent sets of basic data
Initiating event:
Incorrect basic data were made available for clinical dose calculations.
Consequences:
Some patients received a 10% lower dose than prescribed.
Lessons learned:
Make sure you have procedures for not allowing two different sets of data to exist at the same time.

## Slide 33
Treatment planning and dose calculation
9. Wedge factors used twice in calculation of treatment times
A physicist began working in a new institution which had same type of TPS as in his previous work place.
In the planning system of the new institution, wedge factors were already included in the computer calculations. This was not the case in his old institution, where wedge factors were applied manually for each patient.
The physicist began applying the wedge factor manually for patients, after the TPS had done so, which meant it was applied twice in calculations.
Notes: There are echoes of the UK accident from the previous part of the workshop in this case.

## Slide 34
Treatment planning and dose calculation
9. Wedge factors used twice in calculation of treatment times
Initiating event:
Incorrect way of calculating treatment time was used.
Consequences:
A patient received a 53% overdose for a boost (wedged) field.
Lessons learned:
It is important to understand how the TPS works.
Check computer calculations manually.

## Slide 35
?
Treatment planning and dose calculation
10. Incorrect calculation using the inverse square law
The prescribed SSD for a patient was 70 cm instead of the usual 80 cm on the Cobalt unit.
The physicist who calculated the dose used an incorrect inverse square correction factor.
Calculations were not checked until after the eighth fraction, when the mistake was discovered.
Notes: It can be emphasized that it is easy to make the mistake of inverting an inverse square correction.

## Slide 36
?
Treatment planning and dose calculation
10. Incorrect calculation using the inverse square law
Initiating event:
Wrong way of distance correction was used.
Consequences:
The patient received 3.4 Gy per fraction instead of the intended 2.0 Gy per fraction.
Lessons learned:
Maintain awareness for unusual treatments.
Calculations should be checked independently.
Don’t allow many treatment-fractions to take place before you check the calculations.

## Slide 37
Treatment planning and dose calculation
Also remember lectures on:
USA, 1987-88
(Computer file not updated for 60Co source change)
UK, 1982-1990
(Lack of procedures for acceptance of a treatment planning system)
Panamá, 2000
(Problems with data entry to a treatment planning computer)
France, 2004-2005
(Erroneous calculation for soft wedges)
UK, 2006
(Incorrect manual parameter transfer)

## Slide 38
Simulation
Case 11
Notes: It is possible to point out that this case can be seen to be fairly trivial and most likely happen from time to time in most clinics.

## Slide 39
Simulation
11. Incorrect labelling of simulator film
A treatment simulation was performed in prone position instead of routine supine position.
The right side of the simulator film was mistakenly marked as being the left side.
The patient was then set up incorrectly (i.e. set up according to the simulator film) on the treatment unit and irradiated to the right side instead of the intended left side.

## Slide 40
Simulation
11. Incorrect labelling of simulator film
Initiating event:
Simulator film was labelled incorrectly
Consequences:
The patient received more than 2 Gy to healthy tissue.
Lessons learned:
Check orientation of the anatomical site relative to the film carefully, not the least when the treatment is simulated in an unusual position.

## Slide 41
Also remember lecture on:
USA, 2007
(Reversal of images)
Simulation

## Slide 42
Treatment set-up and delivery
Cases 12 - 13

## Slide 43
Treatment set-up and delivery
12. Incorrect identification of patient
A radiation technologist called a patient’s name. Another patient responded. The photo in the patient-record was not consulted.
Freckles on the other patient’s back were mistaken for treatment positioning tattoos, while the patient indicated that the set-up was not correct. An oncology physician was called.
The physician verified that the treatment was correct according to the chart, but did not speak to or examine the patient. Irradiation went ahead.
Notes: It is possible to point out that although this type of event is not totally uncommon in clinical reality, it is nevertheless considered to be of such a magnitude that the International BSS deems it to be mandatory to report to the regulatory authorities.

## Slide 44
Treatment set-up and delivery
12. Incorrect identification of patient
Initiating event:
A patient responded when another patient’s name was called.
Consequences:
The patient received 2.5 Gy to healthy tissue (spine).
Lessons learned:
Check patient’s photograph.
Confirm anatomical marks for beam location.
Make sure to follow up if a patient is warning and objecting to being treated at the wrong site.

## Slide 45
ten in twenty
twenty in ten
Treatment set-up and delivery
13. Misunderstanding of a complex treatment plan given verbally
A patient was prescribed a Cobalt treatment to two different treatment sites. Site One: 2.4 Gy per fraction for 20 fractions. Site Two: 2.5 Gy per fraction for 10 fractions.
The two technologists misunderstood the physician’s verbal instructions, in particular in relation to differences in number of treatment fractions.
Therefore, the second site received an additional four days of treatment before the error was detected.

## Slide 46
ten in twenty
twenty in ten
Treatment set-up and delivery
13. Misunderstanding of a complex treatment plan given verbally
Initiating event:
Misunderstanding of verbal instruction.
Consequences:
The patient received an overdose of 40% to one site.
Lessons learned:
Use written procedures for treatment prescription.
Maintain awareness for complex treatments.

## Slide 47
Case histories and specific lessons learned for brachytherapy
Generic lessons learned
Are there recurring themes in the lessons learned?
Next:
Notes: This concludes the cases from external beam therapy.

## Slide 48
References
INTERNATIONAL COMMISSION ON RADIOLOGICAL PROTECTION. Prevention of Accidental Exposures to Patients Undergoing Radiation Therapy. ICRP Publication 86, Volume 30 No.3 2000, Pergamon, Elsevier, Oxford (2000)
INTERNATIONAL ATOMIC ENERGY AGENCY. Lessons learned from accidents in radiotherapy, Safety Reports Series No. 17, IAEA, Vienna (2000).