## Slide 1: Module 2.9: HDR unit malfunction (USA)
IAEA Training Course

## Slide 2: HDR remote afterloader
A small and mobile source housing installed in shielded treatment room
Remote control console
Source is 192Ir, with apparent activity of 4.3 Ci (160 GBq)
Source is attached to a wire
  can be extended under remote control through one or more catheters in succession into the patient
A second wire with a dummy source
  used first to verify the pathway through the catheter and to verify positions etc…
Example of HDR unit
Notes: <footer>
Older HDR units, especially those installed in developing countries, used cesium or cobalt sources. Iridium is preferred because it has a higher specific activity and therefore can be made smaller. The source used in the device involved in this accident is less than 1 mm in diameter, and about 1 cm long.
The disadvantage of using iridium is its short half-life. The source must be replace three or four times each year.
The various HDR units available today use slightly different technologies to assure that the catheters are properly connected to the HDR unit and that they are not twisted or kinked in a way that might cause the active source to get caught or be exposed unintentionally. However, they all use a dummy source that is extended into each catheter first, before the active source is extended.

## Slide 3: Background
16 Nov. 1992: Elderly patient being treated for anal carcinoma at Indiana Regional Cancer Center (IRCC)
  The patient was scheduled for 3 treatments of 6 Gy each
    Omnitron 2000 HDR unit
  Five catheters were placed into the target volume
Typical treatment
Notes: <footer>
Delivery of multiple treatment sessions is customary with HDR. In this case, the patient had the catheters implanted on 16 November, and was scheduled to receive the first of three treatments of 6 Gy that day. The remaining two treatments were to be given at one-week intervals.
Treatment is essentially automatic. The HDR unit was connected to the five catheters, the staff left the treatment room, and started the treatment from the console. The HDR unit guided the dummy source into the first catheter to ensure its patency, then guided the active source into that catheter. At the end of the planned treatment time, the source was removed from the first catheter. Then the dummy source entered the second catheter, returned to the safe, and the active source was positioned in the second catheter for the prescribed treatment time. This process was repeated for the third and fourth catheters.
The dummy source successfully navigated the fifth catheter, but when the active source was directed into the fifth catheter, an error was reported and the source returned to the safe. The staff attempted to restart the treatment in the fifth catheter several times, but the unit reported an error each time and returned the source to the safe.

## Slide 4: What happened?
The dummy source was introduced without any problems
With the HDR source
  Four channels went well
  Upon attempting to direct the source into the fifth catheter, the control console reported an error
  After several attempts, the treatment was abandoned

## Slide 5: What happened?
Termination of the treatment
The staff entered the treatment room
  disconnected the HDR unit from the implanted catheters
  removed the patient
A typical treatment setup
Notes: <footer>
It seems inconceivable that the staff could ignore the area radiation alarm. However, the staff at this center reported frequent problems with the alarm, and had become accustomed to ignoring it. In fact, it appears that during this accident, they unplugged it because the alarm was annoying.
It also seems inconceivable that they didn’t use a survey meter, which was available to them, to survey the patient. However, they apparently relied upon the console which indicated that the source had returned to the safe.
A problem with the console design was that the console messages were cryptic and limited. In fact, the console was only reporting that the source wire had been fully retracted.

## Slide 6: What happened?
An area radiation alarm indicated high radiation levels, but was ignored
Both sound and sign alarm
The staff reported that the alarm “often malfunctioned” and were used to ignore it
Typical room monitor system

## Slide 7: What happened?
A survey meter was available but was not used to confirm or rule out the area alarm’s signal
The HDR console reported that the source was “safe”
The patient was transported back to her nursing home
Typical hand held meters

## Slide 8: The accident
The hospital staff did not recognize that the source had broken loose from the guide wire, and had remained inside the catheter
The catheters remained in the patient, with the HDR source, as the patient was transported back to the nursing home
20 Nov. 1992 (4 days later) - the catheter containing the source fell out
  The catheter (and source) was placed in a red “medical biohazard” trash bag
  Later, the bag was moved to another storage location with other trash bags where it remained until 25 Nov.
Notes: <footer>
The source was actually in the patient for nearly four full days. At this time, the activity of the source was approximately 4.3 Ci (159 GBq).
When the catheter fell out, the catheter and source were placed in storage with other trash for another five days.
During this time (from 16 to 25 November) the nursing home staff were exposed to the source.

## Slide 9: The discovery
25 Nov. 1992: A driver from the waste handler picked up the red-bag bio-waste from the nursing home
The package was loaded it into a truck trailer with other trash, and transported to the company’s facility in Carnegie, PA.
From there it was transported to a facility in Warren, OH.
Notes: <footer>
The error might have been discovered earlier, but IRCC had not used their HDR unit since the patient was treated on 16 Nov.

## Slide 10: The discovery
A radiation detector at the Warren facility identified radiation emissions from the trailer, and the facility ordered the trailer to return to Carnegie
The trailer remained at the Carnegie facility until Monday, 30 Nov. 1992
Typical vehicle meters

## Slide 11: The discovery
1 Dec. 1992: The trailer was searched, the bag containing the source was found, and traced back to the nursing home
The nursing home was contacted, and they in turn notified the hospital
At this point, the source had been missing for 16 days without notice
The medical physicist confirmed that the HDR source was missing
  The unit had not been used since the event!

## Slide 12: Cause of the accident
The source in this type of HDR unit was welded to the guide
The source was shipped to the Cancer Center in a shielded cask that employed Teflon near the source
In the presence of moisture, radiolysis produced hydrogen fluoride which reacted with the Nitinol1 wire, corroding the weld
The corrosion ultimately weakened the weld, and the source broke off the wire when stressed
weld
source
guidewire
Typical emergency container
1) Nitinol is an alloy of nickel and titanium.
Notes: <footer>
The manufacturer has changed the design of the shipping cask to remove teflon.
The manufacturer also changed the design of the source wire to reduce the risk of separation of the source from the wire.

## Slide 13: Another event
A similar event occurred at another facility three weeks later (Pittsburgh, PA)
The Pittsburgh physicist was aware of the previous event and responded quickly, avoiding significant dose consequences to the patient or staff
Notes: This highlights the importance of sharing information on accidents and also learning from events in other clinics.

## Slide 14: Dose estimates
Patient:
  Prescribed dose was 18 Gy in 3 fractions
  Actual dose delivered over four days at the prescription point (1 cm from catheter) was 16 000 Gy (16 kGy, 1,600,000 rads)
  1000× over dosage
  An NRC consultant concluded that the radiation overexposure likely was the cause, or at least a contributing factor, of the patient’s death
  She died one day after the catheters fell out
Notes: <footer>
The dose to the patient was estimated from the source activity and the time during which the source was lodged inside the patient (almost four days). The dose prescription point was at 1 cm from the catheters, so this point was used to estimate the total dose received.
The doses received by other patients, visitors, workers and hospital staff were estimated from the time they believed they were in the vicinity of the patient, and their distance from the patient.
All together, the NRC’s consultants evaluated the doses received by 94 individuals.
No occupational worker received a whole body radiation dose above the NRC occupational limit (for one calendar quarter) of 12.5 mSv.
Numerous members of the public received doses exceeding the applicable limits, but none received a dose from which acute radiation injury or significant sequelae would be expected.
No contamination of individuals or property occurred.

## Slide 15: Dose estimates
Nursing home staff:
  Several staff were in close contact as they cared for the patient during these four days
  One nursing assistant received doses between 0.7 Sv and 1.6 Sv to the hands
Hospital staff:
  Doses were relatively low, because they were in contact with the patient only briefly

## Slide 16: Dose estimates
Members of the public:
  Various individuals were exposed including
    Patients
    Waste disposal workers
    Patient transporters
    Visitors to the nursing home
  Doses to these individuals were estimated to range between 0.4 mSv and 220 mSv
A total of 94 people were irradiated
  9 people have not been identified

## Slide 17: Actions after 2nd Source Disconnected
The NRC and the United States Food and Drug Administration (FDA) requested to either discontinue using the Omnitron 2000 Unit or immediately implement the following actions
  In accordance with 10 CFR 35.404(a), the licensee shall make a radiation survey of the patient with an appropriate radiation detection survey instrument to confirm that all sources have been removed. This survey is in addition to any indication of radiation levels provided by an area radiation monitor. The surveys shall be done immediately after completion of the therapy procedure prior to removal of the patient from the shielded HDR treatment room, and appropriately documented with initials/signatures.
  A licensee shall not conduct any procedure for which a decoupled source cannot be removed expeditiously from the patient and placed in a shielded condition. The licensee shall have written emergency procedures describing actions to be taken, including surgical intervention, should the source not return to the shielded container at the conclusion of treatment. The licensee shall assure that appropriate staff and equipment are available immediately, at the location that the HDR procedure is performed, to implement the written emergency procedures. Equipment should include shielded storage containers, remote handling tools, and, if appropriate, supplies to help surgically remove sources from the patient to include scissors and cable cutters. The emergency source removal procedure should minimize exposure to healthcare personnel while maximizing safety to the patient.
  The licensee shall ensure that personnel are trained in both the routine use of the device and emergency procedures to return the source to a safe condition. Training shall be provided immediately for new personnel and retraining provided semiannually for all personnel. The licensee shall retain the records of this training for a period of three years.
From NRC BULLETIN 92-03 (Dec issue)
Notes: This is an advanced bullet list where only the highlighted text is supposed to be mentioned.

## Slide 18: FDA 1993 May 19
The methods used in, and the facilities and controls used for, the article's manufacture, packing, storage, and installation are not in conformity with applicable requirements under the current good manufacturing practice regulations.

## Slide 19: Lessons: Radiotherapy Department
Ensure that all staff
  Are properly trained in radiation safety procedures
  Are properly trained in the operation of equipment
  Are properly trained for emergency situations
Include in the Quality Assurance Program
  Formal procedures for verifying the proper operation of the HDR remote afterloading equipment before patient treatments
  Formal procedures for verifying the operation of radiation safety equipment
  Formal procedures for using radiation safety equipment when radioactive materials are used for therapy
Routine surveys of HDR patients to ensure that the source has returned properly to the shield after treatment
Procedures mandating the use of personal dosimeters by staff

## Slide 20: Follow up - Biloxi, MS - 1994
A patient was prescribed a lung brachytherapy treatment delivered by an Omnitron 2000 high-dose- rate (HDR) remote afterloader system.
The prescribed tumor treatment plan included 10 Gy absorbed doses at five stop positions, using a 144.3 GBq 192Ir source
At the end of the fifth and last treatment, the source wire retracted 0.5 cm and stopped and got stuck
Alarms immediately alerted the staff that the source remained inside the patient's body
They followed emergency procedures, entered the therapy room, and removed the catheter from the patient
Once outside the patient's body, the radioactive source retracted to the stored position
U.S. Nuclear Regulatory Commission
Notes: This slide can be excluded

## Slide 21: Reference
Nuclear Regulatory Commission. Report to Congress on Abnormal Occurrences. 92-18. Loss of Iridium-192 Source and Medical Therapy Misadministration at Indiana Regional Cancer Center in Indiana, Pennsylvania. NUREG-0090, Volume 15, No. 4. US NRC, Washington DC (1992)