## Slide 1: Module 1.2: BSS appendix 2 (medical exposure)
IAEA Training Course
Notes: The main focus of this workshop is on accidental exposures in medical exposures, and appendix 2 of the International BSS deals with medical exposures, which will be gone through in more detail in this lecture.

## Slide 2
BSS Appendix II
BSS Appendix II addresses:
Responsibilities of registrants, licensees, and medical practitioners
Justification of medical exposures
Optimization of protection for medical exposures
Guidance levels for diagnostic exposures
Dose constraints for volunteers
Maximum patient activity at time of discharge
Investigation of accidental medical exposures and incidents
Maintenance of records
Notes: These are the different sections of appendix 2, which will be gone through one-by-one.

## Slide 3
BSS Appendix II
BSS Appendix II addresses:
Responsibilities of registrants, licensees, and medical practitioners
Justification of medical exposures
Optimization of protection for medical exposures
Guidance levels for diagnostic exposures
Dose constraints for volunteers
Maximum patient activity at time of discharge
Investigation of accidental medical exposures and incidents
Maintenance of records
Notes: Starting with responsibilities.

## Slide 4
Responsibilities
Registrants and Licensees shall ensure that:
Exposures are given only when prescribed by a medical practitioner
Overall patient protection be assigned to medical practitioners
Sufficient number of trained medical and paramedical personnel are available

## Slide 5
Responsibilities
Registrants and Licensees shall ensure that:
For radiotherapy, the calibration, dosimetry and quality assurance be performed by or under the supervision of a qualified expert in radiotherapy physics
Training criteria be specified or be subject to approval, as appropriate, by the Regulatory Authority in consultation with relevant professional bodies
Notes: It is of great importance in radiotherapy to ensure that a qualified medical physicist is closely involved in the dosimetric and geometric integrity of the whole treatment chain.

## Slide 6
Responsibilities
Comment: In radiotherapy practice …
Radiotherapy is multidisciplinary in nature - a variety of professionals are involved.
This includes professional groups such as medical physicists, radiotherapy technicians, dosimetrists and engineers specialized in radiotherapy equipment as well as oncology and patient support staff such as oncology nurses, dieticians, social workers and other allied health professionals.
In regard to all physical aspects of the radiation delivery the medical physicist (BSS: “qualified expert in radiotherapy physics”) carries the ultimate responsibility
Notes: Many groups are involved in radiotherapy, but it should be emphasized again that the local medical physicist carries the responsibility for the physical aspects of the radiation delivery. It will be examplified in case studies what might happen when this is not clear.

## Slide 7
Responsibilities
Medical practitioners shall promptly inform the registrant or licensee of any deficiencies or needs regarding compliance with the Standards with respect to protection and safety of patients and shall take such actions as may be appropriate to ensure the protection and safety of patients

## Slide 8
BSS Appendix II
BSS Appendix II addresses:
Responsibilities of registrants, licensees, and medical practitioners
Justification of medical exposures
Optimization of protection for medical exposures
Guidance levels for diagnostic exposures
Dose constraints for volunteers
Maximum patient activity at time of discharge
Investigation of accidental medical exposures and incidents
Maintenance of records

## Slide 9
Justification
Justification of medical exposure:
The benefit has to balance the risk, taking into account the benefits and risks of available alternative techniques that do not involve medical exposure
Guidelines (e.g. WHO) should be considered

## Slide 10
Justification
Justification of medical exposure:
Radiological exams should have clinical indications, provide information on health, or be otherwise justified
Screening is justified only if the expected advantages for the individuals examined or for the population as a whole are sufficient to compensate for the economic and social costs, including the radiation detriment

## Slide 11
Justification
Justification of medical exposure:
Exposure of humans for medical research shall be:
  in accordance with the Helsinki Convention
  evaluated by an Ethical Review Committee
Examination as part of theft detection purposes is deemed not justified, but if conducted is considered non-medical exposure
Notes: For theft-detection (e.g. screening at a diamond mine), the argument might be thought to be a bit “circular”.

## Slide 12
BSS Appendix II
BSS Appendix II addresses:
Responsibilities of registrants, licensees, and medical practitioners
Justification of medical exposures
Optimization of protection for medical exposures
Guidance levels for diagnostic exposures
Dose constraints for volunteers
Maximum patient activity at time of discharge
Investigation of accidental medical exposures and incidents
Maintenance of records
Notes: Optimization is of great importance in the context of accidental exposures, and the subsections will be more closely viewed.

## Slide 13
Optimization
Subsections of “Optimization of protection” deal with requirements addressing:
  Design of equipment and facilities
  Operation of facilities for diagnosis and treatment
  Calibration of sources and equipment
  Clinical dosimetry
  Quality assurance

## Slide 14
Optimization 1: Design
Sources for medical exposure shall meet requirements elsewhere in BSS, and shall be designed so that:
failure of single component is detected promptly to reduce risk of unintended exposure
likelihood of human error in the delivery of unplanned medical exposure is minimized

## Slide 15
Optimization 1: Design
Registrants and licensees shall:
use information from supplier to anticipate errors
take reasonable measures to prevent errors
  employ qualified personnel
  establish procedures for calibration, QA, protection
  provide proper training and regular retraining
take measures to reduce consequences of errors
develop contingency plans to respond to errors, display plans, and conduct drills

## Slide 16
Optimization 1: Design
Registrants and licensees in collaboration with suppliers shall ensure in relation to equipment (sources):
conformation to ISO / IEC standards
There should be awareness in developing countries that they have the power to require compliance by suppliers and to make use of acceptance tests, which are universally shared by manufacturers

## Slide 17
Optimization 1: Design
Registrants and licensees in collaboration with suppliers shall ensure in relation to equipment (sources):
performance specifications, operating, maintenance and safety instructions provided in a major world language understandable to the user and translated into local languages when appropriate
console displays parameters and instructions in a major world language acceptable to the user

## Slide 18
Optimization 1: Design
Registrants and licensees in collaboration with suppliers shall ensure in relation to equipment (sources):
beam control mechanism be fail-safe and clearly displayed
exposure be limited to area being examined
radiation field be as uniform as practical
exposure rates outside area being examined be kept as low as reasonably achievable

## Slide 19
Optimization 1: Design
Registrants and licensees in collaboration with suppliers shall ensure in relation to diagnostic equipment (sources):
generator design facilitates keeping of medical exposures as low as reasonably achievable consistent with obtaining adequate diagnostic information
operating parameters are clearly and accurately displayed
exposures are terminated automatically (time / mAs / dose)
fluoroscopy requires constant pressure on a switch, and indications are provided of elapsed time or surface dose

## Slide 20
Optimization 1: Design
Registrants and licensees in collaboration with suppliers shall ensure in relation to therapeutic equipment (sources):
operating parameters can be selected and are displayed clearly, accurately and unambiguously
sealed sources return to safe in event of failure
high-energy equipment:
  have two independent fail-to-safety mechanisms to terminate irradiation
  allow operation only in conditions selected at control panel

## Slide 21
Optimization 1: Design
Registrants and licensees in collaboration with suppliers shall ensure in relation to therapeutic equipment (sources):
bypassing of interlocks can be performed only by maintenance personnel using proper keys or codes
tele- and brachytherapy sources comply with requirements for “sealed sources”
monitoring equipment gives warning of unusual situation

## Slide 22
Optimization 2: Operation
Registrants and licensees shall ensure in relation to diagnostic exposure operation that medical practitioners who prescribe or conduct exams:
ensure that appropriate equipment is used
ensure that the exposure of patients be the minimum necessary to achieve the required diagnostic objective, taking into account norms of acceptable image quality established by appropriate professional bodies and relevant guidance levels for medical exposure
consider previous exams to avoid unnecessary exposure

## Slide 23
Optimization 2: Operation
Registrants and licensees shall ensure in relation to diagnostic exposure operation that:
technologists select the relevant parameters so that their combination produces minimum patient exposure fulfilling the clinical purpose of the examination
portable/mobile equipment be used only when necessary
examinations that might expose an embryo or fetus be avoided if possible
shielding of radiosensitive organs be provided when appropriate

## Slide 24
Optimization 2: Operation
Registrants and licensees shall ensure in relation to therapeutic exposure operation that:
exposure of normal tissue during radiotherapy be kept as low as reasonably achievable consistent with delivering the required dose to the planning target volume, and organ shielding be used when feasible and appropriate
radiotherapeutic procedures causing exposure of the abdomen or pelvis of women who are pregnant or likely to be pregnant be avoided unless there are strong clinical indications
patients be informed of possible risks
Notes: The last point is perhaps difficult to achieve. Information on risks need to be put in a context and weighted against potential benefits.

## Slide 25
Optimization 3: Calibration
Registrants and licensees shall ensure that:
the calibration of sources used for medical exposure shall be traceable to a Standards Dosimetry Laboratory
the calibration of radiotherapy equipment shall be in conformance with a protocol (such as IAEA TRS-277 based on air kerma standards)*
*at the moment of issuing the BSS the TRS 398 (based on absorbed-dose-to water standards) was not yet issued

## Slide 26
Optimization 3: Calibration
Registrants and licensees shall ensure that:
sealed sources be calibrated in terms of activity, RAKR, or absorbed dose: medium, distance and date specified
unsealed sources be calibrated in terms of activity at time of use
calibration be part of commissioning of treatment unit, following service work, and at intervals approved by Regulatory Authority

## Slide 27
Optimization 4: Clinical Dosimetry
Registrants and licensees shall ensure these items are determined and documented:
radiological examinations: representative dose values
external beam radiotherapy: min. and max. dose to PTV as well as dose to relevant point (center) and other points deemed significant
brachytherapy: absorbed dose at relevant points
treatment with unsealed sources: representative dose
all radiotherapeutic treatments: dose to relevant organs
Notes: Documentation of absorbed dose in external beam radiotherapy follows the recommendations of ICRU (Reports 50 and 62).

## Slide 28
Optimization 4: Clinical Dosimetry
Registrants and licensees shall also ensure that, within acceptable ranges of good practice and equipment performance:
the prescribed dose at the prescribed beam quality is delivered to the PTV
doses to other tissues and organs are minimized

## Slide 29
Optimization 5: Quality Assurance
Registrants and licensees shall establish a comprehensive QA programme which shall include:
measurements of physical parameters of radiation generators, imaging devices and irradiation installations at commissioning and periodically thereafter
verification of appropriate physical and clinical factors used in patient diagnosis and treatment
Notes: If there are no national guidelines for commissioning and quality control of radiotherapy equipment, it could be of value to look towards the recommendations from international professional bodies, such as AAPM or ESTRO.

## Slide 30
Optimization 5: Quality Assurance
Registrants and licensees shall establish a comprehensive QA programme which shall include:
written records of relevant procedures and results
verification of calibration and operating conditions of dosimetric and monitoring equipment
as far as possible, regular and independent quality audit reviews of the QA programme for RT procedures

## Slide 31
BSS Appendix II
BSS Appendix II addresses:
Responsibilities of registrants, licensees, and medical practitioners
Justification of medical exposures
Optimization of protection for medical exposures
Guidance levels for diagnostic exposures
Dose constraints for volunteers
Maximum patient activity at time of discharge
Investigation of accidental medical exposures and incidents
Maintenance of records

## Slide 32
Guidance Levels
Registrants and licensees should ensure guidance levels are determined and implemented so that:
corrective actions be taken if doses fall substantially below guidance levels and the exposure does not yield the intended benefit of exam or treatment
reviews be considered if doses exceed guidance levels as an input to ensuring optimized protection of patients

## Slide 33
Guidance Levels
For diagnostic and nuclear medicine exams, guidance levels should be derived from surveys of broad spectrum of exams, and consider entrance doses, beam areas, and radionuclide activities
If such surveys are not available, refer to guidance levels in BSS Schedule III for a guide to performance of equipment

## Slide 34
BSS Appendix II
BSS Appendix II addresses:
Responsibilities of registrants, licensees, and medical practitioners
Justification of medical exposures
Optimization of protection for medical exposures
Guidance levels for diagnostic exposures
Dose constraints for volunteers
Maximum patient activity at time of discharge
Investigation of accidental medical exposures and incidents
Maintenance of records

## Slide 35
Dose Constraints
When individuals are to be exposed for research purposes, without likelihood of direct benefit, the Ethical Review Committee shall specify dose constraints to be followed.

## Slide 36
Dose Constraints
Registrants and licensees shall constrain any dose to individuals incurred knowingly while voluntarily helping (other than in their occupation) in the care, support or comfort of patients undergoing medical diagnosis or treatment, and to visitors to patients who have received therapeutic amounts of radionuclides or who are being treated with brachytherapy sources, to a level not exceeding that specified in Schedule II of the BSS.

## Slide 37
BSS Appendix II
BSS Appendix II addresses:
Responsibilities of registrants, licensees, and medical practitioners
Justification of medical exposures
Optimization of protection for medical exposures
Guidance levels for diagnostic exposures
Dose constraints for volunteers
Maximum patient activity at time of discharge
Investigation of accidental medical exposures and incidents
Maintenance of records

## Slide 38
Max patient activity at discharge
Patients who have received therapeutic procedures with sealed or unsealed radionuclides shall not be discharged until the activity within the body falls below the level specified in Schedule III of the BSS.

## Slide 39
BSS Appendix II
BSS Appendix II addresses:
Responsibilities of registrants, licensees, and medical practitioners
Justification of medical exposures
Optimization of protection for medical exposures
Guidance levels for diagnostic exposures
Dose constraints for volunteers
Maximum patient activity at time of discharge
Investigation of accidental medical exposures and incidents
Maintenance of records

## Slide 40
Investigation of accidental medical exposures
Registrants and licensees shall promptly investigate:
therapeutic treatment to wrong patient, wrong tissue, with wrong pharmaceutical, or with dose or dose fractionation differing substantially from the values prescribed by the medical practitioner or which may lead to undue acute secondary effects
diagnostic exposure significantly greater than intended or repeated so as to exceed guidance levels
equipment failure, accident, error or mishap with potential for causing patient exposure significantly different from that intended
Notes: In the first point, the phrase “differing substantially” is not specific in terms of dose deviation, but will have to be looked at individually. In the last point, potential erros are also highlighted as indications for investigation under certain circumstances.

## Slide 41
Investigation of accidental medical exposures
If investigation is required, registrants shall:
calculate or estimate dose and dose distribution received
determine corrective measures to prevent recurrence
implement corrective measures under their responsibility
submit a written report to Regulatory Authority describing the above
inform the patient and his doctor about the incident

## Slide 42
Investigation of accidental medical exposures
Non-optimized procedures in interventional radiology can also cause accidental exposures
In this case, the arm should not be in the beam for a cardiac procedure.
This was preventable by providing a restful position to the arm

## Slide 43
BSS Appendix II
BSS Appendix II addresses:
Responsibilities of registrants, licensees, and medical practitioners
Justification of medical exposures
Optimization of protection for medical exposures
Guidance levels for diagnostic exposures
Dose constraints for volunteers
Maximum patient activity at time of discharge
Investigation of accidental medical exposures and incidents
Maintenance of records

## Slide 44
Records
Registrants and licensees shall maintain and provide when requested by the Regulatory Authority, the following records:
diagnostic radiology: information needed to make retrospective dose assessment from exposures
nuclear medicine: radiopharmaceuticals administered and their activities
exposure of volunteers in medical research
Also shall have available records of calibration and QA checks

## Slide 45
Records
Registrants and licensees shall maintain and provide when requested by the Regulatory Authority, the following records:
radiotherapy: description of PTV, dose to center of PTV, maximum and minimum doses to PTV, doses to other relevant organs, dose fractionation and overall treatment time