## Slide 1: PREVENTION OF ACCIDENTAL EXPOSURE IN RADIOTHERAPY
Part 2: Case histories of
major accidental exposures in radiotherapy
IAEA Training Course

## Slide 2: Overview / Objectives
Module 2.1: Incorrect decay data (USA)
Module 2.2: Erroneous use of TPS (UK)
Module 2.3: Accelerator software problems (USA and Canada)
Module 2.4: Computer file not updated (USA)
Module 2.5: Incorrect repair of accelerator (Spain)
Module 2.6: Miscalibration of beam (Costa Rica)
Module 2.7: Errors in TPS data entry (Panama)
Module 2.8: Accelerator interlock failure (Poland)
Module 2.9: HDR unit malfunction (USA)
Module 2.10: Accident update – some newer events (UK, USA, France)

## Slide 3
Overview / Objectives
Group exercise G2: Case studies
Group exercise G3: Cobalt source stick
Objectives:
To review and analyze case histories of major accidental exposures, including both over- and under-exposures
To derive specific lessons learned from each of these case histories

## Slide 4: Module 2.1: Incorrect decay data (USA)
IAEA Training Course

## Slide 5
Background
A cobalt unit was used for teletherapy at Riverside Hospital in Columbus, Ohio, USA
This unit was initially calibrated correctly
Cobalt unit (not the actual unit in Ohio)

## Slide 6: Background
During the period 1974-1976 the physicist failed to perform regular measurements (calibrations and QA)
The physicist relied on estimations of the decay of the source to predict dose rate and calculate treatment time
Rather than calculated decay, the physicist plotted dose rate on graph paper and extrapolated

## Slide 7
What happened?
Decay was determined from straight-line plot on semi-log graph paper with calendar ordinate

## Slide 8
What happened?
When edge of graph paper was reached, physicist continued plot on linear paper

## Slide 9
The physicist used a continuation page that had linear scales on both axes
This created two problems:
  Linear Y-axis did not correspond to log Y-axis, so straight line extrapolation resulted in ever-more incorrect output values
  Linear X-axis did not correspond to calendar axis, so extrapolation led to incorrect date values
What happened?

## Slide 10
These errors in predicting the dose-rate were made by the physicist in the time period 1974-1976
The errors resulted in the dose-rate being under-estimated by 10% to 45%
The patients received corresponding overdoses of 10% to 55%
Magnitude of error increased almost linearly with time
Magnitude of accident

## Slide 11
Magnitude of accident
Aug-74
Nov-74
Mar-75
Jun-75
Sep-75
Jan-76
Apr-76
10
100
Year/Month
Patient Overdoses
Percent Overdose [%]
50

## Slide 12
Discovery / investigation of accident
The incident came to light because patients started exhibiting symptoms of overexposure
The accident was investigated by the US Nuclear Regulatory Commission
Notes: It is very important to also be aware of unusual reactions to treatment and to report and investigate any of these incidents immediately.

## Slide 13
Investigation: further complications
When requested, the physicist produced ten calibration documents showing the correct machine output
  These were discovered to have been fabricated
  The output of the cobalt unit had not been checked for 22 months
Notes: This demonstrates the importance of independent secondary checks and carried out by a professional from a different discipline where possible.

## Slide 14
Impact of accident
426 patients received significant overdoses
11 were untraced - 415 followed up
795 sites at risk identified
57% (243) died within the first year
In 87 patients there was local control with no documented recurrence
Survivors beyond the second year had an increased frequency of complications
Notes: Overdosage has a significant impact on mortality and morbidity and all actions have consequences.
Following this incident there was a high incidence of late effects in normal tissue

## Slide 15
Impact of accident
426 patients received significant overdoses
Notes: This is a graphical representation of the survival data over a fifteen year follow up period.

## Slide 16
Impact of accident
74 patients (17%) had significant complications within the first year post-treatment
78 patients survived more than 5 years
  68%: tumour free
  63%: no complications

## Slide 17
Site-specific complications
22 patients had skin complications
51 patients had bladder and rectal complications
15 patients had respiratory complications
6 patients had cardiovascular complications
18 patients had nervous system complications
Total number of complications was 154 of 795 sites (19%)
Notes: Symptoms related to site treated but this slide shows the wide range of significant complications recorded and the percentage of patients affected.

## Slide 18
During the follow up period 1-3 years after treatment, a high proportion of the patients developed significant, often life-threatening, complications
  Severe skin reactions with ulceration
  Mucosal reactions with necrosis
  Stenosis of pharynx and oesophagus
  Ulceration and/or perforation of stomach and bowel
  Bone necrosis and myelopathy
Site-specific complications

## Slide 19
Follow-up and utilization of knowledge
Follow-up charts were available for between 1 and 3 years after treatment
Follow up period of several decades
Notes: This incident did highlight the importance of good data and record keeping. In this hospital follow up in the original charts was available for the first three years and the follow up period for the patients is now over 20 years.

## Slide 20
To increase the body of knowledge on tolerance limits of normal tissue
To improve thereby the efficacy and safety of curative radiotherapy
Accident allowed an analogous clinical cell ‘experiment’ on human subjects
Follow-up and utilization of knowledge
Notes: This information was analysed scientifically and used to increase the knowledge on tolerance limits of normal tissue. While obviously not carried out deliberately, this incident allowed an analogous clinical cell “experiment” on human subjects.
Given the practice at the time to treat a “one field on alternate days” technique the results showed the increased risks associated with the technique and provided data on the specific effects of larger fractions for a given total dose and overall time
Sharing of the information with the wider community allowed others to learn

## Slide 21
Database for analysis
In relation to the onset of early and late reactions in specific organs and tissues at various times after exposure
  Dose
  Fractionation
  Time
  Volume
Notes: The database included information on dose, fractionation, time and volume which could then be related to the onset of early and late reactions in specific organs and tissues at various times after exposure.

## Slide 22
Lessons: Radiotherapy Department
Include in the Quality Assurance Programme:
  Provide independent check of physicist’s work
  Formal procedures for calibrating treatment unit on a regular schedule
  Department should provide sufficient staff to handle workload
  Records must accurately document performance of accepted QA procedures
  Establish an accurate database for follow-up

## Slide 23
Lessons: Radiotherapy Department
In case of unusual reactions in one patient notified by a technologist or directly by the patient, the radiation oncologist should immediately request the medical physicist to perform a verification to detect a possible error in any of the treatment steps
Unusual reactions in more than one patient should lead to a request to the medical physicist to immediately verify the dosimetry of the treatment unit

## Slide 24
References
Cohen L, Schultheiss T E, Kennaugh R C. A radiation overdose incident: initial data. Int J Radiat Oncol Biol Phys 33: 217-224 (1995)
ICRP Publication 86: Prevention of accidental exposures to patients undergoing radiation therapy (2000)