## Slide 1: Module 3.3: Incidents in any clinic
IAEA Training Course

## Slide 2
Accidents versus incidents
Accident:
Any unintended event, including operating errors, equipment failures and other mishaps, the consequences or potential consequences of which are not negligible from the point of view of protection or safety.
Incident:
Any unintended event, including operating errors, equipment failures, initiating events, accident precursors, near misses or other mishaps, or unauthorized act, malicious or non-malicious, the consequences or potential consequences of which are not negligible from the point of view of protection or safety.
(Source: IAEA Safety Glossary, 2007)
Notes: This is one way of defining an accident versus an incident. An incident is often seen to be of a smaller magnitude than an accident.

## Slide 3
Accidents versus incidents
Accidents:
e.g.
The nine cases of major accidental exposures presented in modules 2.1 – 2.9
Many of the cases presented in modules 2.10, 3.1 and 3.2
Notes: We have so far mainly presented accidents in the workshop, except for some of the more minor events in modules 3.1 and 3.2.

## Slide 4
Accidents versus incidents
Incidents:
e.g.
Some of the cases presented in modules 3.1 and 3.2
The events presented in this module 3.3
Notes: More of these minor events will be presented here.

## Slide 5
ICRU 62 - “... a dose difference as small as 5% may lead to real impairment or enhancement of tumour response, as well as to an alteration of the risk of morbidity.”
Incidents are important
Notes: There are several reasons why incidents should be considered when addressing the safety in radiotherapy. Firstly, one should not disregards incidents that affect the intended dose to the patient with a few percent. Follow this example of TCP / NTCP curves 5% up or down to see how a dose deviation affects the probablity for tumour control or normal tissue complication negatively.

## Slide 6
Variable magnitude:
Many incidents (e.g. mistake in calculation of monitor units for a single patient) can have a variable magnitude (e.g. for Patient 1, the mistake causes a dose deviation of 5%, while for Patient 2, the same type of mistake causes a dose deviation of 50%).
Incidents are important
Notes: Another reason to consider incidents

## Slide 7
More events:
Incidents are more numerous than accidents, so there are more opportunities to learn and improve the safety, than by only looking at major accidents.
Incidents are important
Notes: A third reason to consider incidents

## Slide 8
Actual incident:
The unforeseen event has affected the treatment of the patient
Potential incident:
“Near miss” - The unforeseen event was discovered and halted before it affected the treatment of the patient
Incidents
Notes: Incidents can be divided into “actual” and “potenital”, like this.

## Slide 9
In this module:
Data from a clinic, on incidents originating from events in treatment planning and calculation, are presented and analysed
This clinic is well-equipped and well-staffed – i.e. “if it happens here, it can happen anywhere”
Incidents
Notes: Actual data from St. Luke’s Hospital, Dublin, are presented in the remainder of this module.

## Slide 10
Clinical environment
Around 4500 new patients per year
Six linear accelerators
One 3D treatment planning system
Notes: First, the clinical environment of this radiotherapy clinic is presented, in order to put the incident-data in a context for the participants to relate to.

## Slide 11
Clinical environment
Notes: The number of computer calculated treatment plans per year has been rising steadily with a doubling-time of nine years. (Other plans are manually calculated, I.e. single fields or parallel-and-opposed fieds). This indicates a clinical environment where complexity is growing.

## Slide 12
Clinical environment
Safety system for treatment planning in the clinic:
Many check stations to ensure the “quality of the output” from treatment planning
Notes: This flowchart gives a picture of the “check-stations” (yellow boxes) that a plan passes through when going from preparation to being used for treatment. The red arrows indicate that results from check-stations are passed back to earlier parts of the process (forming closed loops).

## Slide 13
Clinical environment
Safety system for treatment planning in the clinic:
Incident data presented in this lecture: found before treatment through primary and secondary calculation checks (potential incidents) or through weekly overview checks or vigilance during treatment (actual incidents)
Notes: The incident data presented comes from three types of plan-checks: primary (first check by physicist before treatment); secondary (second check by radiographer on treatment unit before treatment); and weekly overview (done by radiographer on treatment unit while treatment is in progress).

## Slide 14
Clinical incident data
Categories in IAEA database of radiation accidents:
Equipment design
Calibration of beams
Maintenance
Treatment planning and dose calculation
Simulation
Treatment set-up and delivery
Notes: When comparing with the categories in modules 3.1 and 3.2, the incidents covered in this module 3.3 only come from the “Treatment planning and dose calculation” category.

## Slide 15
Overall:
Data from five years of checking treatment plans and calculations
Data from nearly 28000 plans / calculations:
  Manual plan calculations (calculating monitor units or treatment time without planning system - TPS)
  Computer plan calculations (TPS-based calculations)
Clinical incident data
Notes: Results from a great amont of checks / calcualtions form the basis for the data of incidents.

## Slide 16
How many incidents originate in treatment planning?
In ~3 % of all plans, primary checking found an unintended “potential incident”
In ~½ % of all plans, secondary checking (after primary) found an unintended “potential incident”
Actual incidents in ~¼ % of cases
For each actual incident, ~14 potential incidents were found through calculation checking
Clinical incident data
Notes: Mistakes were found in many plans. Usually by the first check, sometimes with the second check, but some went through all checks and became actual incidents.

## Slide 17
What type of incidents originate in treatment planning?
In ~2.6 % of simple manual plans, there was a potential incident
In ~4.3 % of the more complex computer plans, there was a potential incident
Clinical incident data
Notes: It is interesting to see that the more advanced computer planning showed a higher frequency of potential incidents.

## Slide 18
Types of mistakes made (in 17503 manual plans)?
Clinical incident data
Notes: The following four slides show different parameters that had a mistake done with them. In this case, all mistakes could be seen as originating in the act of manually transferring information (e.g. looking up a value in a table, writing in a value in a chart, etc). This slide is valid for manual plans. The column to the left is the parameter in question, to the right is the number of events relating to this parameter. Far right at the bottom is the overall percentage.

## Slide 19
Types of mistakes made (in 17503 manual plans)?
Clinical incident data
Notes: Manual plans. Incidents relating to creating a parameter or manually calculating a parameter.

## Slide 20
Types of mistakes made (in 10327 computer plans)?
Clinical incident data
Notes: Transfer-incidents for computer plans.

## Slide 21
Types of mistakes made (in 10327 computer plans)?
Clinical incident data
Notes: Incidents in creating / calculating for computer plans.

## Slide 22
Clinical incident data
Overall:
There was a “potential incident” in planning originating in the act of manual transfer of information in 1.4% of plans
There was a “potential incident” in planning originating in the act of creating or calculating of information in 1.8% of plans

## Slide 23
Summary
An incident frequency of 3% could be seen in a “normal clinic”. Most of these potential incidents were stopped before they became actual incidents (14 : 1) through a good safety system.
TCP tells us that an incident with a few percent impact can have a negative impact on the intended treatment outcome.
Many incidents have a variable magnitude: the next time, the same incident could become an accident.

## Slide 24
Summary
Incidents are more numerous and varying than major accidental exposures …
…so make sure you learn from the incidents happening in your clinic, to avoid an accident!

## Slide 25
References
Holmberg O. Ensuring the intended volume is given the intended absorbed dose in radiotherapy - Managing geometric variations and treatment hazards (ISBN 91-628-6002-X) (2004)
Holmberg O, McClean B. Preventing treatment errors in radiotherapy by identifying and evaluating near misses and actual incidents. J Rad Ther Practice 3:13-25 (2002)