## Slide 1: Module 2.3: Accelerator software problems (USA and Canada)
IAEA Training Course

## Slide 2: Background
Image supplied by
William Van de Riet
Mid 1970s - AECL developed a new double-pass concept for electron acceleration
  needs less space to develop similar energy levels
  dual-mode linear accelerator
  more compact and versatile than the older Therac-20
Therac 25 took advantage of computer’s abilities to control and monitor hardware
Notes: This project was in corporation with the French company CGR.

## Slide 3: Background
From Aldrich 1985 (top) and O’Brien 1985 (bottom)
Notes: At the top you can see the new wave guide with the electron gun (light blue) to the right close to the bending magnet (green). The electrons are injected into the standing wave accelerator tube and accelerated to the left against the gantry. The RF is fed into the guide structure at the middle. Reaching the gantry they exit the waveguide and are transported through a 180 degree bending magnet (pink) and the re-entering the tube at the correct phase. The electrons are accelerated a second time reaching a maximum energy of 25 MeV at the entrance to the achromatic 270 degree bending magnet that directs the electrons against the isocenter (or patient).
More information can be found in the Aldrich and O’Brien papers.

## Slide 4: Photon vs. electron treatment head
From O’Brien 1985
Notes: This construction is originally from the CGR machine where the electron field is not formed by a fixed application as is common today but by trimmers attached to the photon collimator blocks (jaws). One may aslo note that there is now scattering foils for electrons. Instead the flatness of the beam is achieved by scanning the electron beam over the field opening.

## Slide 5: Filter carousel
Notes: For this dual modality machine a filter carousel is used to support the different modes as well as the field light. Thus, for electrons the scanning magnet is used, for photons the flattening filter is rotated into position, and when the patient is set-up the light source and the mirror is in position. All this is controlled by the software together with mechanical switches indication the position of the carousel. There is also a mechanical plunger to improve the positioning.

## Slide 6: Scanned electron beam
From O’Brien 1985
Notes: Two different scanning patterns are used for the electron beam. A forth-and-back pattern is used for high energies and circular-type is used for the lower energies.

## Slide 7: What happened?
Notes: While looking at one of the promotion products from the company we will now continue with the scenario and the chain of events that happened in the mid 1980-ies.

## Slide 8: Event 1 - Marietta
1
Notes: Let’s go to Marietta, Georgia, USA

## Slide 9: 1 - Marietta, June 1985
Illustration of chest treatment with electrons (Nucletron)
Approximately 6 months experience with the new machine
A breast cancer patient treated with 10 MeV electrons commented You burned me! after the radiation session
The treated area felt warm when the technologist checked

## Slide 10: 1 - Marietta, June 1985
The physicist was informed
He contacted AECL and asked if the machine could deliver electrons without scanning
3 days later – AECL responded – impossible!
No report of misadministration (not required in those days) was filed

## Slide 11: 1 - Marietta, June 1985
The patient’s shoulder became red and swollen
The oncologist judged this either as:
  disease related complication, or
  normal complication during radiotherapy
Two weeks later the physicist noticed:
  a red area on the exit side of the treated field
  skin started to peel off
  the shoulder was immobile
  patient felt a great pain

## Slide 12: 1 - Marietta, June 1985
Neither operators nor AECL could believe it was the linac
This accident was not recognized or reported
The manufacturer did not:
  document this problem
  describe the event to other users or the FDA
  take any further action
When looking back – The printout function was disabled and no permanent record of the treatment existed
Notes: This accident was reported later when the other accidents became known.
The company who was contacted by the physicist did not file this, communicate with other users etc…

## Slide 13: Event 2 - Hamilton
Notes: One month later in Hamilton, Ontario, Canada

## Slide 14: 2 - Hamilton, July 1985
About 6 months’ use of the linac
A cervix patient was treated
During the 24th fraction a H-Tilt error message showed up and terminated the treatment
Dosimetry system indicated
  no dose
  treatment pause
The operator continued by simply pressing key “P” to proceed
  this was permitted five times by the software
The operators were familiar with errors that inhibited the treatment with no consequences for the patients
The malfunctions simply reported a numeric code, with no explanation in documentation, etc.

## Slide 15: 2 - Hamilton, July 1985
A technician was called
  who found nothing wrong with the accelerator
After the treatment the patient complained
  burning sensation in her hip
  felt like an electrical shock
  three days later: still burning, hip pain and excessive swelling

## Slide 16: 2 - Hamilton, July 1985
The linac was taken out of clinical use a suspected overdose
AECL was informed and sent an engineer to the site
The local authorities were informed

## Slide 17: 2 - Hamilton, July 1985
Actions by AECL
  The malfunction could not be reproduced
  The switches indicating the carousel position were assumed to be the error source
  This was confirmed when switches were intentionally disabled and showed that ambiguous position could be interpreted by the software
  Revised software, etc., were released
  In final report, ”…cannot be firm on the exact cause of the accident but can only suspect…”
  Safety was improved by 5 magnitudes, it was claimed!
Notes: Comment: A 5 time magnitude of improved safety is 100,000 times safer product. What did the company compare with? Did they have any fault checking system? Actually no-one knows where this figure comes from.

## Slide 18: 2 - Hamilton, July 1985
Message to users from company
  Check position of carousel
  Discontinue treatment if H-tilt pops up combined with erroneous dose
  However, no information about a patient injury
Suggestion from a user at a user meeting to add a hardware interlock for carousel position – ignored by company
Patient died from her disease Nov. 1985
  Autopsy showed:
        severe radiation damage
        hip replacement indicated
Notes: This time the company released

## Slide 19: Event 3 - Yakima
Notes: Let’s move west to Yakima, Washington, USA

## Slide 20: 3 – Yakima, Dec. 1985
The linac was modified, according to the vendor, in Sep. 1985
Female patient developed erythema during X ray treatment in December 1985
Parallel reddish striped pattern over the hip
Reaction was considered normal during EBT course
After completing treatment in Jan./Feb. the skin reaction was considered abnormal and further investigated

## Slide 21: 3 – Yakima, Dec. 1985
Illustration of shadow-tray
Source of striped pattern?
  Open slots in shadow-tray possible source
    tray and skin pattern did not match
  Chemotherapy reaction was ruled out
    the other ports did not show any reaction
  Patient slept with heating bag
    after X ray examination: pattern of wires in heating bag did not match
Notes: The slide showing the scanning pattern could be shown in context with this one instead of in the introduction of the design.

## Slide 22: 3 – Yakima, Dec. 1985
Contact with AECL resulted in: ”…we are of the opinion that this damage could not have been produced by any malfunction of the Therac-25 or by any operator error.”
Followed by two pages outlining: ”…apparently been no other instances of similar damage to this or other patients.”
The skin reaction was classified and filed as ”Cause unknown”

## Slide 23: 3 – Yakima, Dec. 1985
The staff were not aware that this was the 3rd incident
They were even told by the vendor that none had occurred
The company should by now be aware of three suspicious events
Notes: Incident or accidental exposure or overdosage

## Slide 24: Event 4 - Tyler
Notes: Then it is time for Tyler, Texas, USA

## Slide 25: 4 – Tyler, March 1986
Linac had been operational for two years, >500 patients treated
Male patient was to receive his 9th fraction with 1.8 Gy to his back
  22 MeV electrons

## Slide 26: 4 – Tyler, March 1986
Operator was very familiar with accelerator console, and typed quickly and efficiently
She noticed that she had selected “X” for X rays instead of “E” for electrons
She moved the cursor up to correct the entry
She then hit the return key several times to skip to the bottom of the screen, and pressed “B” for “beam-on”
After a moment, the console displayed ‘Malfunction 54’ and ‘treatment pause’

## Slide 27: The screen layout of Therac-25

## Slide 28: 4 – Tyler, March 1986
According to documentation at the console, this was a “dose input 2” error
Only 6 monitor units had been delivered of 202 requested, so the operator pressed “P” to proceed
The accelerator shut down again

## Slide 29: 4 – Tyler, March 1986
On this day, the video and audio monitors were disconnected due to a fault
  The operator was not aware of what happened in the treatment room
  The patient “had got himself up from the treatment table, and was banging on the treatment room door to be let out”
The operator quickly opened the door

## Slide 30: 4 – Tyler, March 1986
The patient had been irradiated while moving on the treatment table during the second attempt
The patient was exposed twice and reported impressions:
  1st exposure
    Electrical shock or as “hot coffee was poured over his back”
    Heard a buzzing sound from the linac (abnormal)
  2nd exposure
    The arm was hit by an electrical shock
    The hand was leaving his body
Notes: Since the patient had have several fractions prior to this one the patient was rather familiar with the sound during treatment.

## Slide 31: 4 – Tyler, March 1986
The oncologist examined the patient, concluding:
  the intense erythema resulted from an electric shock
The medical physicist investigated and found:
  nothing unusual in the operation of the accelerator
  the calibration of the machine was correct
Illustration of calibration
Notes: <footer>

## Slide 32: 4 – Tyler, March 1986
No one understood the error message and treatments continued the rest of the day
The manufacturer’s service staff were called in, and arrived the next day
  They found no problems – testing the linac a full day
  When asked, they said “there had been no reports of accidents involving radiation overexposure with this equipment”

## Slide 33: 4 – Tyler, March 1986
AECL response
  Concluded it was an electrical shock
  However, they should have been aware of the three suspicious events
An external consultant checked the grounding of the machine – it was found to be OK
After re-checking calibration, the linac was in clinical use in April again
Notes: <footer>

## Slide 34: Event 5 - Tyler
Notes: <footer>
Still in Texas

## Slide 35: 5 – Tyler, April 1986
A second event in the same clinic occurred three weeks later
A patient was scheduled to have a 10 MeV electron field on one side of his head
Same operator, same X-to-E change on keyboard
After beam on - a loud noise was heard through the intercom system
Illustration of head-and-neck treatment
Notes: <footer>

## Slide 36: 5 – Tyler, April 1986
Illustration of head-and-neck treatment
A Malfunction 54 error was displayed and the machine paused
The operator went into the room, where the patient was on the treatment table - moaning

## Slide 37: 5 – Tyler, April 1986
The patient complained about feeling of “fire” on his cheek
The operator called the physicist immediately
The patient described the event as “something hit me on the face”, flash of light, and heard a sizzling sound as when frying eggs…
Notes: <footer>

## Slide 38: 5 – Tyler, April 1986
Three weeks later the patient died with disorientation which finally led to coma, high fever, neurological damage, etc.
Autopsy showed that the temporal lobe and the brain stem had massive acute radiation injuries

## Slide 39: 5 – Tyler, April 1986
The physicist took the machine out of clinical service
  He contacted the manufacturer to report this second problem
He and the operator then started their own investigation
  After some effort they were able to recreate the sequence of events
  They concluded that the speed of entering commands was the key factor
  Measurements by the physicist showed
    the monitor chamber was saturated
    the dose was somewhere above 40 Gy
Notes: Here we now have our heroes in this process of events

## Slide 40: 5 – Tyler, April 1986
The manufacturer was not able to reproduce the malfunction until the physicist explained that the commands had to be typed quite quickly
  Two days later AECL came back with report
    Measured dose could run up to 250 Gy
    The “frying sound” was electrical discharges in the ion chamber

## Slide 41: Investigation in Tyler
It was shown that the malfunction was permitted by the design of the software
A subroutine monitored the entry of treatment parameters, including mode, energy, collimator setting, etc.
When all entries were completed, and the cursor returned to the command line at bottom, a flag was set
Notes: Now we are into some software discussion and why the software did not recognise the editing on the console.

## Slide 42: Investigation in Tyler
Once the flag was set indicating all entries were completed, if an entry was changed and the cursor returned quickly to the command line, the subroutine would not detect that a change had been made
Another subroutine controlled the position of the carousel according to the mode selected
The operator was able to correct the mode from “X” to “E”, triggering rotation of the carousel, but without properly setting the electron current or the scanning magnets

## Slide 43: More information
AECL person testified later
  A cursor up problem had occurred earlier (Marietta) in 1985
  Another cursor up problem at a non-mentioned site had occurred in the summer of 1985 (Hamilton?)
  These should have been fixed
  There would be no relationship between these errors and the Tyler incidents!
Review of software at a Therac 20 site by a user
  The typing errors (when students use the linac) makes a lot of fuses blow!
  The same X/E editing problem exists but the hardware interlock prohibits beam-on
Notes: <footer>

## Slide 44: Why high patient dose?
High beam current for photons ~ 100-1000 x
No beam scanning → flattening filter
Extremely high electron energy fluence
Measurements by the physicist at Tyler initially reported only doses of ~8 Gy
  Ion recombination/discharges/saturation in the ion chamber was introducing a large error, and re-evaluated dose was found to be ~40Gy
Manufacturer determined dose to be 250 Gy
The dose varies from site to site due to tuning of linac (number of radiation pulsed within 1st 3 s)
Notes: <footer>
The efficiency producing photons is in the order of per cent thus the beam current is somewhere between 100 to 1000 times higher when x-ray therapy is given compared to electrons. Further on, when X is selected no scanning is used, a flattening filer should be in the beam. These two parameters together will give an extremely high dose rate if electrons are given while the software think it is photons. First the high beam current and secondly the beam is not scanned even if the correct position of the carousel is used.

## Slide 45: Response from authorities and vendor
FDA response - after report from Texas Health Department (Tyler)
  The Therac-25 linac defective
  The company was ordered to inform all purchasers, investigate the problem, and submit a corrective plan (FDA must approve)
AECL informed users about a temporary fix to allow treatment
  The up key may not be used
  It must be removed from the keyboard to avoid usage and the contactors taped
  They did not mention what the dangers were!
FDA did not approve this modification
A revised Corrective Action Plan (CAP) was presented by AECL which was approved by FDA
Notes: <footer>

## Slide 46: Event 6 - Yakima
Notes: <footer>
It is now about 8 month after the last accident in Tyler, Texas. The next event will be on the west coast again, i.e. Yakima in Washington.

## Slide 47: 6 – Yakima, January 1987
A patient was to have a port film taken with 4 and 3 MU plus a 79 MU treatment
Open field
35 x 35 cm2
for orientation
with 3 MU
Treatment field
18 x 22 cm2
with 4 + 79 MU
Notes: <footer>
A double exposure for verification was prescribed.

## Slide 48: 6 – Yakima, January 1987
First the 4 MU to the treatment field was given
The collimator opened automatically and the second exposure with 3 MU was given
The operator entered the room to verify the patient position
He checked the beam position relative to the skin marks by turning on the field light (the carousel rotated)
Notes: <footer>
You have to remember here that field light can only be used after the carousel has rotated the mirror/light system into position. Then we also have the filters out of position.

## Slide 49: 6 – Yakima, January 1987
The operator reset the machine for treatment,
  went out of the room and
  accidentally forgot the film under the patient
Beam-on was started by pressing key “B”
The display showed no dose or dose rate
After 5-6 s the unit shut down with a quickly disappearing message
The machine was paused, thus the operator pressed “P” to proceed
Notes: <footer>

## Slide 50: 6 – Yakima, January 1987
The machine started again, but terminated with a flatness message
The operator could hear the patient saying something and entered the room
The patient had felt a burning sensation in the chest

## Slide 51: 6 – Yakima, January 1987
The console displayed 7 MU (the two exposures for the film)
Later the same day the patient developed skin burn over the treatment area
Four days later, it was striped (similar to the previous problem at Yakima)

## Slide 52: 6 – Yakima, January 1987
AECL response
  All tests indicating perfect performance
  All users were informed to check the carousel position visually
  Suspected that the light source had been in position instead of the photon filter
The patient film was compared with tests which confirmed the erroneous position
Notes: <footer>

## Slide 53: 6 – Yakima, January 1987
It was concluded that a dose of 40 to 50 Gy could have been given per irradiation, i.e. up to 100 Gy to the specific patient instead of 0.86 Gy (two film exposures + treatment)
According to AECL the corrections (not yet implemented) resulting from the Tyler accidents would have prevented this

## Slide 54: Consequences of the over dosage
1 | The breast was removed, complete loss of arm and shoulder mobility and constant pain
2 | Patient severely burned, died November 1985
3 | Tissue necrosis, chronic ulcerations and pain in the treated area and several skin grafts were made and the patient is alive today (1995)
4 | Pain in neck and shoulders, periodic nausea and vomiting, radiation induced myelitis of the cervical cord, paralysis of the left arm and both legs, left vocal cord and left diaphragm Died after five months due to the complications
5 | Patient died one month later
6 | The patient died within weeks from a severe terminal form of cancer but the over-dosage probably shortened his life
Notes: 1) June 1985, Marietta, Georgia, USA:
  Estimated dose > 80 Gy probably 150-200 Gy (local physicist)
2) 26 June 1985, Hamilton, Ontario, Canada:
  Estimated dose 130-170 Gy (by an AECL employee)
3) December 1985, Yakima, Washington, USA:
  When the staff was aware after their second incident, they called the patient
4) 21 March 1986, Tyler, Texas, USA
5) 11 April 1986, Tyler, Texas, USA
6) 17 January 1987, Yakima, Washington, USA

## Slide 55: Characteristics of the accidents
Three cases involved carousel rotation prior to treatment (confirmed)
The accelerator malfunctioned shortly after “beam on”, reporting a malfunction code at the console
  The codes were cryptic and not recognized by the operator as indicating a serious error
In several cases, the operator repeated the exposure one or more times
Following treatment, the patients complained of burning sensations, sometimes accompanied by a feeling of electric shock
In each case, the patients received doses of between 40 and 250 Gy in a very brief exposure (1-3 seconds)
Notes: <footer>

## Slide 56: Summary of causes of accidental exposure
Manufacturer recycled software
  Earlier model functioned somewhat differently, so software was not entirely suitable
  Newer model relied entirely on software for safety, whereas older model had mechanical and electrical interlocks
  The safety of the newer system was not evaluated as a whole, only the hardware was evaluated since software had been in use for years…
The manufacturer had no mechanism for investigating and reporting accidents
  After the first accident, the manufacturer refused to believe the equipment was at fault
  The FDA was not notified, nor were other users
  The vendor kept their opinion that this machine was safe
Notes: <footer>

## Slide 57: February 1987
FDA
  …requests the AECL to immediately notify all purchasers and recommend that use of the device on patients for routine therapy be discontinued…
Canadian government
  …AECL to recommend all users in Canada to discontinue the operation of Therac-25…
AECL July
  Informed all users that FDA have approved their latest CAP (about 30 different changes)
Notes: <footer>

## Slide 58: Time line of events
June 1985
Marietta
Jan. 1987
Yakima
June 1985
Hamilton
Dec. 1985
Yakima
Mar. 1986
Tyler
Apr. 1986
Tyler
Feb. 1987
FDA request
July 1987
AECL action plan
Notes: This is a timeline to show the period under which these events occurred. Not all steps are shown here but the accidents and the final actions are included.
The last item is that the FDA approved the Corrective Action Plan from AECL that covers about 30 items.

## Slide 59: Lessons: Radiotherapy Department
Listen, observe, report, and follow-up patient reactions
Investigate all reports of abnormal operation
Include in the Quality Assurance Program
  Review of procedures for reporting unusual events
Do not accept vendor’s answer
  It is always a local responsibility
Notes: <footer>

## Slide 60: Lessons: Suggested by computer science consultants
Documentation key from beginning
  Use established software engineering practices
  Keep designs simple
  Build in software error logging & audit trails
Extensive software testing and formal analysis at all levels
  Revalidate re-used software
Don’t rely only on software for safety
Do incorporate redundancy
Pay careful attention to human factors
Involve users at all phases
Notes: <footer>

## Slide 61: The continuation
Image supplied by Larry Watts
The 11 machines were refitted with the safety devices required by the FDA and remained in service
No more accidents were reported from these machines
After the Therac-25 deaths, the FDA made a number of adjustments to its policies in an attempt to address the breakdowns in communication and product approval.
  In 1990, health care facilities were required by law to report incidents to both the manufacturer and the FDA.

## Slide 62: References
O’Brien et al. Radiation protection aspects of a new high-energy linear accelerator. Med Phys 101-107 (1985)
Aldrich et al. Characteristics of the photon beam from a new 25 MV linear accelerator. Med Phys 619-624 (1985)
O’Brien et al. Characteristics of electron beams from a new 25 MeV linear accelerator. Med Phys 799-805 (1985)
Leveson N G, Turner C S. An investigation of the Therac-25 accidents. IEEE Computer 26: 18-41 (1993)