## Slide 1: Module 2.4: Computer file not updated (USA)
IAEA Training Course

## Slide 2: Background
Section of typical cobalt source
Sacred Heart Hospital, Cumberland, MD
March 1987 – an old 60Co source was exchanged
All corresponding data was updated by a consulting medical physicist…

## Slide 3: What happened?
Typical trimmer bars
…except data for treatment with trimmer bars
The oncologist stated to the consulting physicist that trimmer bars would not be used for treatment anymore
The data file/program for this treatment type was not updated by the physicist

## Slide 4: What happened?
September 1987 – The oncologist started treatment with trimmer bars for whole brain
The staff used the old computer file for calculation of treatment time with trimmer bars
Because the file contained the outdated source activity, patient treatment times were too long

## Slide 5: Discovery of accident
Therapy staff observed skin erythema on several patients during the time period September 1987 and October 1988
They occasionally expressed their concern to the hospital oncologist
  Judged as normal during radiotherapy
October 1988 - the consulting medical physicist was notified
  Found a non-updated computer data file or program

## Slide 6: Decay of 60Co
Activity
Time correction
Source replacement
Erroneous correction
1.75x
Notes: This graph shows how the source disintegrate and how the activity is decreasing (light blue line). To correct for this, the treatment time is increased accordingly (green line).
By not updating the data file with the new source strength one continue to correct the treatment time as if we have a low activity. The correction should start from unity all over at the time of the source replacement.
The time correction built into the data file used for calculation treatment time when trimmer bars is used leads to a larger error as time passes. After 50 month from the original calibration and when the file was updated there is an deviation of 1.75 i.e. 75% overdosage.

## Slide 7: Impact of accident
Patients received doses 75% greater than prescribed
33 patients were treated with this technique during a time period of 13 months
20 patients were dead (either during the course of treatment or after conclusion of treatment) at the time of notification to the authorities

## Slide 8: Actions taken
The computer file was removed
Physicist reviewed all remaining software and data and to perform output measurements
Independent oncologist was hired to review all patient records
Consulting physicist visits hospital weekly

## Slide 9: Lessons: Radiotherapy Department
Develop procedures for indicating clearly software that is commissioned for clinical use,
  and software that has been removed from clinical service
Include in the Quality Assurance Program
  Formal procedures for returning medical equipment to service after maintenance
  Procedures for verifying correct function of software for patient calculations
  Performance of manual calculations to confirm computer calculations of treatment time (and in-vivo dosimetry)

## Slide 10
Reference
Nuclear Regulatory Commission. Report to Congress on Abnormal Occurrences. NUREG-0090, Volume 11, No. 4. US NRC, Washington DC (1988)