## Slide 1: PREVENTION OF ACCIDENTAL EXPOSURE IN RADIOTHERAPY
Part 5: Reporting, investigating and preventing accidental exposures
IAEA Training Course

## Slide 2: Overview / Objectives
Module 5.1: Preventive measures
Module 5.2: Reporting and investigating
Group exercise G5: Reporting incidents
Group exercise G6: Identifying what to do next
Objectives:
To provide knowledge of incident reporting and investigation mechanisms. To apply lessons learned on causes and contributing factors to the prevention of accidental medical exposures

## Slide 3: Module 5.1: Preventive measures
IAEA Training Course

## Slide 4
Generalizing the lessons learned
Working with Awareness and Alertness, Procedures, Training and Understanding, Responsibilities
What can be done to prevent accidental exposures? Preventive measures – an overview
The concept of “Multilayered defence”
Examples of how initiating events could have been prevented from becoming accidental exposures
A checklist for prevention of accidental exposures
Overview of lecture
Notes: The aim of this lecture is to provide a template by which the lessons learnt from the incidents can be applied in practice. Some sample exercises are given towards the end

## Slide 5
Generalizing the lessons learned
Working with Awareness and Alertness Accidental exposures have occurred owing to inattention to details, lack of alertness and lack of awareness. This could also be made worse if personnel have to work in conditions prone to distractions.
Procedures
Accidental exposures have occurred when there is a lack of procedures and checks, or when they are not comprehensive, documented or fully implemented.
Notes: This slide, and the next one, point back to what was learnt in modules 3.1 and 3.2, when generalising the lessons learned. They expand on the four general “headings” that were found and point forward to how this can be best used for prevention.
It is easy to think that once procedures / protocols are in place no accidents can occur. However it is equally important that staff do not become complacent about what they are doing and remain constantly alert and aware to anything that differs from what would normally be expected.

## Slide 6
Generalizing the lessons learned
Training and Understanding
Accidental exposures have occurred when there is a lack of qualified and well-trained staff, with necessary educational background and specialized training.
Responsibilities
Accidental exposures have occurred when there are gaps and ambiguities in functions of personnel and lines of authority and responsibility. Safety critical tasks can be insufficiently covered.
Notes: The education of all professional groups to a high level is essential in radiotherapy. Too often, for instance, the technologist’s level of education is low and does not receive high priority. All professionals involved in the delivery of radiation must have a level of understanding such that they will be in a position to identify potential incidents or deviations from what would be considered the ‘norm’
Clear delegation of tasks and levels of responsibilities must be defined and all involved personnel must understand exactly what is expected of them

## Slide 7
Preventing accidental exposures
Keep these “lessons learned” in mind!
There are preventive measures that can be taken by your clinic to avoid accidental exposures. Some specific measures will be presented for:
- Organization
- Education and training
- Acceptance testing and commissioning
- Follow-up of equipment faults
- Communication

## Slide 8
Preventing accidental exposures
Organization
The structure of a radiotherapy department should be clearly defined, especially with regard to the role and responsibilities of each staff member. Each individual should understand their position in the structure as well as that of others.
Assignment of functions and lines of authority and responsibility should ensure that there are no gaps or ambiguities which leave safety critical tasks insufficiently covered. There should be clear and known job descriptions.
- Example: Who is assigned responsibility for the correct use
of a TPS? (e.g. case history on wedge factors used twice in calculation of treatment times)
Notes: The structure and responsibility of management will be discussed in a separate lecture but there will be some elements of overlap or repetition

## Slide 9
Preventing accidental exposures
Organization cont.
There should be key staff available, and special consideration should be given to the workload, which could become excessive in the case of installation of new equipment, set-up of a new technology or an increase in patient load or complexity of treatment.

## Slide 10
Preventing accidental exposures
Education and training
A critical component of the radiotherapy process is qualified personnel. All staff dealing with radiation sources and patients must have the necessary educational background and specialized training.
Training should be consistent with the responsibilities assigned to each group of professionals, and should include review and analysis of typical accidents.
- Example: A maintenance engineer should not modify parameters on a treatment unit without sufficient knowledge and expertise on that particular type of unit (e.g. case history on incorrect repair followed by insufficient communication – Spain, 1990)
Additional training should be provided when new equipment and techniques are being introduced.
Notes: An initial high level of education for all staff involved in the radiotherapy process is essential. The concept of learning elements of practice together is interesting and helps also to build team skills.
Education should continue throughout the working life of staff to ensure that they remain up to date with current practice.
Continuing education should be resourced and supported from within the organisation and should be mandatory

## Slide 11
Preventing accidental exposures
Acceptance testing and commissioning
Acceptance tests are crucial to the prevention of accidental exposures, because the safety of the system as a whole can be tested, including safety interlocks that can’t be tested under normal operation.
Validation of the equipment and its operation, also under clinical conditions, has to be made before initiating treatment of patients.
- Example: Also perform measurements using other gantry angles than 0°. (e.g. case history on loose wedge mounting mechanism)
.

## Slide 12
Preventing accidental exposures
Acceptance testing and commissioning cont.
Systematic acceptance and commissioning, including a cross-check and independent verification, form a major part of accident prevention. Understanding of equipment should also be considered part of this process.

## Slide 13
Preventing accidental exposures
Follow-up of equipment faults
Lack of follow-up of unexplained malfunctions have ended in fatal accidental exposures.
The most difficult situations are those malfunctions that cannot be reproduced by the maintenance engineers, either due to intermittent faults or faults that only appear in some particular circumstances.
- Example: A fault only occurred when a variable in the accelerator software rolls over zero (i.e. case history on six accidental exposures involving software problems in several accelerators of the same type – Canada and USA, 1985-1987)

## Slide 14
Preventing accidental exposures
Follow-up of equipment faults cont.
A procedure that facilitates interactions between users and manufacturer, a follow-up of unexplained malfunctions, and dissemination of the information to other users of the same type of equipment and to maintenance engineers is a recommended approach.

## Slide 15
Preventing accidental exposures
Communication
There should be clear and concise written rules for communication critical to safety. These rules should be posted and understood.
- Example: Handing over an accelerator to a physicist following maintenance should be formalized and adhered to. (e.g. case history on incorrect repair followed by insufficient communication – Spain, 1990)
Documents critical to safety, for example prescriptions, basic data and treatment plans, should be signed by staff who are responsible and qualified.
Notes: This slide is relevant for all levels of communication and also highlights the importance of correct and comprehensive data collection and recording

## Slide 16
Preventing accidental exposures
Communication cont.
If an unexpected reaction is observed in a single patient, the radiation oncologist should immediately be notified and request the medical physicist to investigate possible errors in the steps of the treatment procedure. If unexpected reactions are observed for multiple patients, the medical physicist should immediately be notified and investigate the dosimetry of the treatment unit.
Notes: This is also an example of the need for awareness and alertness on the part of the radiotherapy technologists in particular who see the patients every day. It also underpins the need for adequate education programmes to ensure that the radiotherapy technologist understands what would be considered to be the normal reaction and an adverse reaction

## Slide 17
Multilayered prevention of accidental exposures
The term “defence in depth” is defined in the BSS as “the application of more than one single protective measure for a given safety objective such that the objective is achieved even if one of the protective measures fail”.
“Defence in depth” can be viewed as several layers of safety provisions, such as physical components and procedures.
Notes: Defence in depth is very much concerned with physical barriers and it is also important to keep in mind the need for conceptual barrier also – i.e. awareness and alertness

## Slide 18
Multilayered prevention of accidental exposures
Multilayered prevention includes aspects of “defence in depth” but also includes aspects such as awareness and alertness which could be termed “conceptual defence”
For this multilayered prevention of accidental exposures to work, these layers need to be independent of each other.
An implemented Quality Assurance program might provide the layers. Part of the QA should be to verify that this is the case!

## Slide 19
Multilayered prevention of accidental exposures
Initiating events will happen many times in any clinic
If there are no layers of safety provision, these events will lead to
accidental exposures

## Slide 20
Multilayered prevention of accidental exposures
Initiating events
Accidental exposures
By putting in a layer of safety- provision, many initiating events are stopped from becoming accidental exposures.
When only a single layer of safety-provision is present, failure of this layer can still lead to accidental exposures.

## Slide 21
Multilayered prevention of accidental exposures
Initiating events
Accidental exposures
By having multiple independent layers of safety-provision, there is a much higher likelihood that accidental exposures are prevented.

## Slide 22
Multilayered prevention of accidental exposures
Examples of multilayered prevention of accidental exposures will be shown for:
        Patient identification error
        Calibration error of treatment unit
        Calculation error in treatment chart

## Slide 23
Multilayered prevention of accidental exposures
Initiating event: Error in patient identification
Consequence: Other patient having healthy tissue irradiated

## Slide 24
Multilayered prevention of accidental exposures
Initiating event: Error in patient identification
Consequence: Other patient having healthy tissue irradiated
Patient-photograph in patient-chart

## Slide 25
Multilayered prevention of accidental exposures
Initiating event: Error in patient identification
Consequence: Other patient having healthy tissue irradiated
Patient-photograph in patient-chart
Procedure to check photography before irradiation

## Slide 26
Multilayered prevention of accidental exposures
Initiating event: Error in patient identification
Consequence: Other patient having healthy tissue irradiated
Patient-photography in patient-chart
Procedure to check photography before irradiation
Procedure to confirm anatomical marks on patient before irradiation

## Slide 27
Multilayered prevention of accidental exposures
Initiating event: Error in patient identification
Consequence: Other patient having healthy tissue irradiated
Patient-photography in patient-chart
Procedure to check photography before irradiation
Procedure to confirm anatomical marks on patient before irradiation
Identity bracelets / identity cards

## Slide 28
Multilayered prevention of accidental exposures
Initiating event: Error in patient identification
Consequence: Other patient having healthy tissue irradiated
Patient-photography in patient-chart
Procedure to check photograph before irradiation
Procedure to confirm anatomical marks on patient before irradiation
Identity bracelets / identity cards
Awareness! React and follow up if the patient indicates that something is wrong

## Slide 29
Multilayered prevention of accidental exposures
Initiating event: Misunderstanding of timer when calibrating Cobalt
Consequence: Wrong dose to all patients treated on Cobalt

## Slide 30
Multilayered prevention of accidental exposures
Initiating event: Misunderstanding of timer when calibrating Cobalt
Consequence: Wrong dose to all patients treated on Cobalt
Independent calibration by other physicist

## Slide 31
Multilayered prevention of accidental exposures
Initiating event: Misunderstanding of timer when calibrating Cobalt
Consequence: Wrong dose to all patients treated on Cobalt
Independent calibration by other physicist
Procedure to compare calibration with source certificate

## Slide 32
Multilayered prevention of accidental exposures
Initiating event: Misunderstanding of timer when calibrating Cobalt
Consequence: Wrong dose to all patients treated on Cobalt
Independent calibration by other physicist
Procedure to compare calibration with source certificate
External postal dosimetry audit

## Slide 33
Multilayered prevention of accidental exposures
Initiating event: Misunderstanding of timer when calibrating Cobalt
Consequence: Wrong dose to all patients treated on Cobalt
Independent calibration by other physicist
Procedure to compare calibration with source certificate
External postal dosimetry audit
In vivo dosimetry

## Slide 34
Multilayered prevention of accidental exposures
Initiating event: Misunderstanding of timer when calibrating Cobalt
Consequence: Wrong dose to all patients treated on Cobalt
Independent calibration by other physicist
Procedure to compare calibration with source certificate
External postal dosimetry audit
In vivo dosimetry
Awareness! There should be shorter treatment times with a new source

## Slide 35
Multilayered prevention of accidental exposures
Initiating event: Mistakenly inverting SSD-correction in MU-calculation
Consequence: Very significant dose deviation for a patient

## Slide 36
Multilayered prevention of accidental exposures
Initiating event: Mistakenly inverting SSD-correction in MU-calculation
Consequence: Very significant dose deviation for a patient
Independent check of calculation

## Slide 37
Multilayered prevention of accidental exposures
Initiating event: Mistakenly inverting SSD-correction in MU-calculation
Consequence: Very significant dose deviation for a patient
Independent check of calculation
Weekly chart-check of “reasonability”

## Slide 38
Multilayered prevention of accidental exposures
Initiating event: Mistakenly inverting SSD-correction in MU-calculation
Consequence: Very significant dose deviation for a patient
Independent check of calculation
Weekly chart-check of “reasonability”
In vivo dosimetry

## Slide 39
Multilayered prevention of accidental exposures
Initiating event: Mistakenly inverting SSD-correction in MU-calculation
Consequence: Very significant dose deviation for a patient
Independent check of calculation
Weekly chart-check of “reasonability”
In vivo dosimetry
Written procedure for calculation methods

## Slide 40
Multilayered prevention of accidental exposures
Initiating event: Mistakenly inverting SSD-correction in MU-calculation
Consequence: Very significant dose deviation for a patient
Independent check of calculation
Weekly chart-check of “reasonability”
In vivo dosimetry
Written procedure for calculation methods
Awareness! Shorter SSD means shorter treatment time for same dose

## Slide 41
Multilayered prevention of accidental exposures
Initiating event: ?
Consequence: ?
TRY IT AS AN EXERCISE!
Examples of initiating events:
Calibration of beam made in penumbra
Pancake chamber used upside down
Use of wedge factor twice in calculation of treatment time
Misunderstanding of verbal prescription

## Slide 42
Checklist for prevention of accidental exposures
Organisation, functions and responsibilities
Have all necessary functions and responsibilities been allocated?
Are all functions and responsibilities understood?
Is the number of staff commensurate to workload?
Is this number re-assessed when workload increases, or when new equipment is purchased?




Notes: It could be suggested to the participants, that this checklist (from ICRP Publication 86) is a good start for “self-assessment” of the clinic.

## Slide 43
Education and training
Is every member of staff educated and trained according to their responsibilities? Is this education and training documented?
Is there a programme for continuing and personal development?
Are lessons from accidents and their prevention included in continued training?
Are there provisions for additional training (new equipment, new procedure)?
Are emergency plans exercises carried out as part of the training?





Checklist for prevention of accidental exposures

## Slide 44
Checklist for prevention of accidental exposures
Acceptance testing and commissioning
Is there a programme for formal acceptance of equipment in place?
Is it carried out according to international or national standards?
Is there a programme of commissioning in place?
Does it include treatment equipment as well as treatment planning systems and simulators and other ancillary equipment?





## Slide 45
Checklist for prevention of accidental exposures
Quality Assurance programme
Is a programme of QA established?
Is the programme based on accepted protocols? Which ones?
Are all tasks of the QA clearly assigned to the right persons?
Are the necessary tools and instruments available?
Are audits part of the programme?






## Slide 46
Checklist for prevention of accidental exposures
Communication
Is a communication policy in place and understood by staff?
Is reporting of unusual equipment behaviour required?
Is reporting of unusual patient reactions required?
Are procedures for equipment transfer for maintenance and return in place?





## Slide 47
Checklist for prevention of accidental exposures
Patient and site identification
Are there procedures to ensure correct identification of patient and site?
Is there a protocol for patient’s chart check?



## Slide 48
Checklist for prevention of accidental exposures
External beam calibration
Are there provisions for initial beam calibration?
Is independent verification in place foreseen and planned?
Is there an accepted protocol? Which one?
Is a programme for follow-up calibration in place?
Is participation in an audit programme part of the programme?






## Slide 49
Checklist for prevention of accidental exposures
External beam treatment planning and in vivo dosimetry
Are treatment planning systems included in the programme of acceptance and testing?
Is treatment planning documented according to accepted protocols?
Are cross-checks and redundant and independent verification included?
Has a system for in vivo dosimetry been considered?





## Slide 50
Conclusions
Keep in mind the lessons learned from accidental exposures in relation to:
Working with Awareness and Alertness
Procedures
Training and Understanding
Responsibilities
Evaluate specific preventive measures that can be taken by your clinic in relation to:
Organization
Education and training
Acceptance testing and commissioning Follow-up of equipment faults Communication

## Slide 51
Conclusions
Initiating events will continue to happen. Take a critical look at the system for multilayered prevention of accidental exposures in your clinic.
Go through the check-list for prevention of accidental exposures in your own clinic, and assess what might be improved.
Despite all measures, accidental exposures may still occur. Be prepared! Have a response plan, with a clear list of actions and responsible officers posted in relevant places. Develop and review this periodically.
Notes: There is a need for continuing vigilance at all times

## Slide 52
References
International Commission on Radiological Protection. Prevention of accidental exposures to patients undergoing radiation therapy. ICRP Publication 86, Volume 30 No.3 2000, Pergamon, Elsevier, Oxford (2000)
International Atomic Energy Agency. Lessons learned from accidents in radiotherapy, Safety Reports Series No. 17, IAEA, Vienna (2000).
rpop.iaea.org