## Slide 1: Module 2.7: Error in TPS data entry (Panama)
IAEA Training Course

## Slide 2: Brief history of the event
Notes: This case study will start with some background and a brief overview of the event itself, befor it goes into a more technical description of what happened.

## Slide 3: Background information
Year 2000, the radiation therapy department of ION was divided between two different hospitals and a total of 1100 patients received radiotherapy.
  Justo Arosemena hospital
  (External beam therapy)
  Gorgas hospital
  (Brachytherapy and hospitalization of in-patients)

## Slide 4
Equipment for external beam therapy (EBT) in Justo Arosemena hospital:
  Cobalt-60 unit (Theratron 780C)
  Cobalt-60 unit (ATC/9 Picker)
  Orthovoltage unit (Siemens Stabilipan)
  TPS (RTP/2 Multidata v.2.11)
  One 60Co unit and the orthovoltage unit were decommissioned and not in use at the time of the accident.
  EBT given from 6 a.m. to 9 p.m. on Theratron (in two shifts).
Background information
Notes: The Picker and Siemens units were not in clinical use.

## Slide 5
Staff of ION:
  Five radiation oncologists
  Two of these radiation oncologists (one in the morning and one in the evening) assigned to Justo Arosemena hospital on a monthly rotation
  Four radiotherapy technologists
  Two medical physicists
  One dosimetrist
Background information

## Slide 6
Factors influencing workload in Justo Arosemena hospital:
  70 to 80 patients treated per day
  Many of these patients treated during the evening with only a single therapist present
  Team divided between two sites
  Multiple fields (SSD set-up technique) with beam modifying devices (blocks and wedges) utilized
Background information
Notes: It is worth to take note of the high workload in this clinic. For example, just one patient with four-fields, SSD set-up on all fields, with blocks and wedges on many fields, would require a lot of effort for a single therapist to treat.

## Slide 7: Brief description of the event
Multidata TPS (2D) used to plan treatment
The TPS allowed four shielding blocks to be entered in any field for calculation of dose distribution
Notes: Here follows a brief account of the event. A more thorough technical discussion will come later in the lecture.

## Slide 8: Brief description of the event
In April 2000 one of the oncologists required one additional block for some treatments in the pelvic region

## Slide 9: Brief description of the event
In order to overcome the limitation of four blocks imposed by the TPS, …
… a new way of entering data was tried (August 2000): to enter several blocks “at once”.
The TPS accepted the data entry, without giving a warning, but calculated incorrect treatment times
Notes: A “home-made” solution to the problem was devised locally – and it seemed to work.

## Slide 10: Discovery of the problem
In November 2000 radiation oncologists observed unusual reactions in some patients (unusually prolonged diarrhoea).
The physicists checked the patient charts but did not find any abnormality (the computer calculations were not questioned)

## Slide 11
Patient charts checked but computer calculations not questioned
Notes: The signatures were all in place (arrows), but computer calculations that formed the basis for the treatment time in the charts were not questioned.

## Slide 12: Discovery of the problem
In February 2001 the error in dose calculations was finally determined
The treatment was simulated on a water phantom and dose measurements were made, which confirmed higher dose
… treatment of relevant patients was suspended.

## Slide 13: The resulting treatment plan
The computer printout provides slightly distorted isodoses but the icon with the blocks was correct.
The treatment time indicated was approximately twice the intended.
Notes: The arrow points towards the icon of the field shape on the treatment plan. It seems that the TPS has taken proper account of the intended field shape, when looking at the icon.

## Slide 14: Remark: findings from quality audits
First audit: February 1999
  Quality controls were made, but written procedures were missing
Second audit February 2001
  Procedures were in place, but no procedure for the use of TPS
  The auditor was not notified of the new approach for data entry
  Tests were performed but not for the specific conditions of this event
Notes: This slide indicates a certain history of quality audits in the clinic.

## Slide 15: Technical description of the problem
Notes: The following section goes into the event in more technical detail.

## Slide 16
The treatment planning system (TPS) at ION:
  Multidata RTD/2
  Version 2.11
  System installed in 1993. Beam data for 60Co entered and verified at this stage.
  This is a 2D TPS. It allows shielding blocks to be entered and taken into account when calculating treatment time and dose distribution.
Treatment planning

## Slide 17
Two of the modules in the Multidata TPS:
  “Dose chart calculator” for calculation of treatment time to a given point
  “External beam” for calculation of treatment time to a given point AND calculation of isodoses
Treatment planning
Notes: It is important that the participants pay attention to the difference between these two modules, or they will not see why events unfolded the way they did.

## Slide 18
Restriction of the treatment planning system:
  Maximum 4 blocks can be digitized for a field in the “External beam” module.
  In the “Dose chart calculator” module, there is no such restriction.
Treatment planning

## Slide 19
Treatments in the pelvic region were performed using “the box technique”.
Up to four blocks per field were often used for these fields.
Treatment planning

## Slide 20: Treatment Planning Entering blocks separately
Menu:
Add 1 block
Type transmission factor
Digitize contour
Repeat the procedure for next block
Notes: <footer>
This slide shows the correct procedure in the TPS for entering a single block. On the left is the wording on the menu on the computer screen.

## Slide 21: Entering four shielding blocks correctly
Notes: Here follows the steps when four shielding blocks are entered in a way that the TPS can correctly handle.

## Slide 23
For some cervix patients, a central shielding was added to the four blocks.
Since no isodoses were requested for these cases, the “Dose chart calculator” module was used. This allows for more than four blocks.
Treatment time was correctly calculated.
Treatment planning

## Slide 24
One of the oncologists started to request isodoses for these patients with five blocks.
The “External beam” module had to be used for this. Because of the four block limitation, initially four or less blocks were digitized.
Treatment time was slightly incorrect due to this. The effect was understood.
Treatment planning

## Slide 25
Staff came up with an approach to enter multiple blocks simultaneously.
This approach was used for fields with four or more blocks. Even though the method was incorrect, the TPS was essentially able to handle this method.
Treatment time was essentially correctly calculated.
Treatment planning
Notes: The staff wanted the TPS to take all blocks into account when calculating the treatment time, so they tried to overcome the “four block limitation” problem. Once they found an approach that seemed to work, it was not only used for five-block fields, but also when fewer blocks were requested, due to the time-saving characteristics of the approach.

## Slide 26: Entering several blocks as one
Notes: Here follows the “home-made” solution for entering four shielding blocks as a single block that was initially devised by staff. Even though this method was wrong, the TPS gave treatment times that did not deviate greatly from the intended. This is “Method 1”.

## Slide 27
Notes: A loop which is not crossing itself is formed with this “Method 1”.

## Slide 28: Variation to new approach
This worked well, but, as the procedure was not written…
…another physicist entered the data in a similar but slightly different way.
This variation causes wrong isodoses and the wrong treatment time.

## Slide 29
Notes: A loop which is crossing itself is formed with this variation: “Method 2”.

## Slide 30: Computer printouts
Notes: Here follows some slides with examples of actual printed out dose distributions from the TPS in Panama.

## Slide 31
Notes: This is a correct dose distribution (open field, no field icon on the side (arrow)). The 50% isodose is also indicated with an arrow.

## Slide 32
Notes: This is a correct dose distribution (four blocks entered correctly, a correct field icon on the side). The 50% isodose is indicated with an arrow.

## Slide 33
Notes: The two previous dose distributions side by side.

## Slide 34
Notes: This is a slightly incorrect dose distribution (four blocks entered with “method 1”, a correct field icon on the side). The 50% isodose is indicated with an arrow. A slight “dent” can be seen on the side (arrow) of the dose distribution, indicating that something is starting to go wrong in the calculations.

## Slide 35
Notes: This is a incorrect dose distribution (four blocks entered with “method 2”, the field icon on the side still seems correct). The 50% isodose is very shallow and the distribution is heavily distorted.

## Slide 36
Notes: Comparison between correct entry (left) and “method 2” (right).

## Slide 37: Second variation – multiple fields
The distortion is not so obvious for a four field treatment.
The icon does not indicate that the TPS is incorrectly used
Calculated treatment time approximately TWICE AS LONG AS INTENDED
Notes: Even though the distorsion of the dose distibution is very clear for a single field, it becomes less obvious in a plan with several fields and a non-flat outline.

## Slide 38
Notes: An actual plan from the Panama TPS, with field shape icons visible on the right side.

## Slide 40
Notes: The effect of “method 2” on TPS calculations of depth dose can be seen here, compared with the depth dose of an open field.

## Slide 41
Calculated treatment time
The calculated treatment time was approximately twice the intended
Example: Treatment time on similar patients had been 0.6 min (one field). Now it had become more than 1.2 min (one field).
Notes: Calculated treatment time with “method 2” became approximately doubled due to the more rapidly falling depth dose shown in the previous slide.

## Slide 42: Discovery of the problem

## Slide 43
In November 2000, radiation oncologists were observing unusually prolonged diarrhoea in some patients.
On request, physicists reviewed charts (double checked). TPS output was not questioned. No anomaly was found.
Discovery of the problem

## Slide 44
In Dec. 2000, similar symptoms were observed. In Feb. 2001, physicists initiated a more thorough search for the cause.
In March 2001, physicists identified a problem with computer calculations. Treatment was suspended.
Nov’00
Dec’00
Jan’01
Feb’01
Mar’01
Symptoms
Chart checks
Symptoms
More thorough
checks
Problem
found
Discovery of the problem

## Slide 45
Isodoses and treatment time were re-examined closer and anomalies were found.
The treatment was simulated on a water phantom and measurement of doses were made, which confirmed higher dose.
Nov’00
Dec’00
Jan’01
Feb’01
Mar’01
Symptoms
Chart checks
Symptoms
More thorough
checks
Problem
found
Discovery of the problem

## Slide 46: Estimation of dose to patients

## Slide 47: Some of the patients treated in the abdominal region were affected In total: 28 patients affected.
Brain: 4.3%
Head and neck: 12.1%
Mamma: 16.8%
Lung: 7.9%
Cervix uteri: 15.5%
Endometrium: 1.5%
Prostate: 9.3%
Rectum: 3.9%
Others: 28.7%
Treatments performed at INO
Notes: This slide gives a background to the diverse treatment given at the clinic in question. Patient groups with potentially affected patients are highlighted.

## Slide 48: Dose estimation
Dose to the 28 affected patients was estimated retrospectively
Dose to prescription point for multiple fields was estimated
Based on the patients’ charts:
  dose rate under reference conditions
  beam set up (depth, effective field, and beam modifiers), and
  treatment times

## Slide 49: Biologically effective dose
Since the dose per fraction was much higher than standard, the biologically effective dose (BED) and the dose equivalent to a treatment of 2 Gy/fraction were also calculated, using the linear quadratic model (α/β = 3 for intestine was used for evaluation of late effects).

## Slide 50: Number of patients and their dose (equivalent to 2 Gy/fraction)
(as of May 30, 2001)
Notes: <footer>
In this figure (illustrating a point in time not very long after the discovery of the accident), the distribution of absorbed dose given to patients who had expired is seen to be shifted towards higher values than the distribution for patients alive at this point in time.

## Slide 51
90
120
150
This accidental exposure
Notes: <footer>
The figure is an indication of the incidence of severe proctitis versus absorbed dose [Gy] to the rectum. The patients affected by this accident are mainly in the interval between the two lines.

## Slide 52
100
This accidental exposure up to 160
Notes: <footer>
The figure is an indication of the incidence of severe large bowel complications versus absorbed dose [Gy]. The patients affected by this accident are mainly to the right of the line in the figure.

## Slide 53
Skin changes even though multiple
fields used
Notes: Even though a four-field technique was used, skin changes could be seen at the entrance of single radiation portals.

## Slide 54
Notes: <footer>
CT image of a patient, showing fibrotic tissue and stenosis of the rectum.

## Slide 55
Notes: <footer>

## Slide 56
Notes: <footer>

## Slide 57
Notes: <footer>

## Slide 58
Dilated air-filled loops of small bowel from a distal small bowel obstruction likely secondary to radiation induced stenosis

## Slide 59: Effects on patients
Effects at the moment of the evaluation mission (May 30, 2001)
8 deaths of 28 patients
5 of these deaths radiation related
2 unknown (not enough data)
1 due to metastatic cancer
20 surviving patients of the affected
Notes: <footer>

## Slide 60: Lessons and recommendations

## Slide 61: Initiating event and contributory factors
The event was triggered by
  The search for a way to overcome the limitation of the TPS (four blocks only)
Contributory factors
  The computer presented the icon as if the blocks were correctly recognized
  The procedure was not tested
  The trick “worked” and was time-saving

## Slide 62: Initiating event and contributory factors
Contributory factors (continued)
  Treatment times were longer than usual but no one detected it
    workload
    limited interaction (radiation oncologists, medical physicists and radiotherapy technologists)
    computer calculations in general were not verified
  Patient reactions were realized but the follow-up was insufficient
Notes: While this slide mentions that no one detected the discrepancies in treatment time, it could alternatively have been the case that something was detected but that there were no channels in which to communicate this.
In the last point, the follow-up is referring to the follow-up of independently verifying that the dose was correct.

## Slide 63: Lessons for manufacturers
Avoid ambiguity in the instructions
Thorough testing of software, also for non-intended use
Guide users with warnings on the screen for incorrect data entry
Be readily available for consultation, especially when a change in the way of use is intended

## Slide 64: Lessons for radiotherapy departments
TPS is a safety critical piece of equipment
Quality control should include TPS
Procedures should be written
Change in procedures should be validated before being put into use
Computer calculation should be verified (manual checks for one point)

## Slide 65: Lessons for radiotherapy departments
Awareness of staff for unusual treatment parameters should be stimulated and trained
Communication should be favoured
Unusual reactions should be completely investigated and dosimetry data tested

## Slide 66
Panama incident summary
‘Minor’ change of practice in use of a treatment planning system
Not systematically verified
8 patients dead

## Slide 67: Reference
IAEA: Investigation of an accidental exposure of radiotherapy patients in Panama (2001)
rpop.iaea.org

## Slide 68
Postscript
Towards the end of 2004, two physicists involved in this event were sentenced to four years in prison respectively, as well as a period of seven years when they were not allowed to practice in the profession.

## Slide 69
Postscript
According to the court, they did not inform their superiors regarding the modifications in practice in relation to the use of the treatment planning software.