## Slide 1: Module 2.6: Miscalibration of beam (Costa Rica)
IAEA Training Course

## Slide 2: Overview
Error in calculation
Overdosage
115 patients affected
Person in charge of dosimetry not appropriately qualified
Failure to act on advice
Notes: The error in this case was compounded by the lack of formal education of personnel involved and the failure to act on prior warning signals

## Slide 3: Prior to the event: problems
The radiotherapy department had participated in the IAEA / WHO Postal Dose Quality Audit
Since 1977 significant differences between the dose stated by the hospital and the dose determined by the IAEA’s laboratory
Since no satisfactory explanation was found, an expert mission was sent in July 1996, shortly before the event
Notes: This slide highlights almost twenty years of documented problems in the clinic. An early response could have identified some of the problems that contributed to the subsequent over exposure
The results of the postal audit are confidential and actions on results of postal audits were / are not binding on the institutions. Practice at the IAEA has also changed and where confidentiality is still maintained an expert is dispatched to carry out a check much earlier following discovery of discrepancy in the TLD postal audit

## Slide 4: Prior to the event: mission
The purpose of the mission was to:
  Find out the reasons for the discrepancy
  Evaluate the physical aspects of Quality Assurance
  Check the application of the IAEA code of practice for absorbed dose determination
  Check compliance with the Basic Safety Standards (BSS)

## Slide 5: Prior to the event: findings
Problems
  The same “output” values were used for PDD as for TAR
  It revealed confusion between the concepts of kerma in air and absorbed dose to water
Action
  The expert brought it to the attention of the radiation oncologists and the person responsible for dosimetry
Notes: <footer>
PDD PERCENTAGE DEPTH DOSE
TAR TISSUE AIR RATIO
This problem should have alerted attention to a lack of understanding on the part of the person responsible for dosimetry in the hospital

## Slide 6: Prior to the event: responses
Responses
  The response was skeptical, arguing that no clinical abnormal reactions had been observed so far
  The expert insisted that the errors detected were underdosage errors, which, would not manifest in acute patient reactions, but would only manifest later in poor tumour control

## Slide 7: Time scale prior to the event
July 1996
Investigation by the IAEA Expert Group
August 1996
The source of the Alcyon teletherapy unit was exchanged and put into operation
Alcyon CGR II unit
Notes: No records were kept on the calibration of radiation beams and there was no information available on the equipment used or the working conditions prevailing
Dosimetry check was carried out one unit in July and confirmed as within acceptable limits TLDs left with person in charge of dosimetry to complete further check following exchange of source

## Slide 8
A radiation oncologist from another hospital (hospital Calderón Guardia) whose patients were treated in San Juan de Dios Hospital (where the event occurred) noticed some unusually severe reactions in some of the patients treated on the Alcyon unit
These reactions were related to the skin and low intestinal track, e.g., diarrhoea and abdominal pain
The event and its discovery
Notes: <footer>
There were three hospital in Costa Rica devoted to the treatment of cancer. Two of the hospitals had radiotherapy equipment. There were 3 Cobalt units in total, two of which were in San Jose and two orthovoltage units
Costa Rica had 5 Radiation Oncologists, 1 Physicist and 9 Radiotherapy Technicians
Of the two hospitals in San Jose hospital 1 (Calderon Guardia) had 1 radiation oncologist but no radiotherapy equipment, the patients were treated in the hospital 2 (San Jan de Dios). Hospital 1 however had better data available
The radiation oncologist in hospital 1 noted the unusual reactions of his patients

## Slide 9
Detection
  He compared the patient charts and realized that the indicated dose rate from the Alcyon was lower than from the Theratron, despite the fact that the activity of the new unit was higher
Action
  He reported it to the person responsible for the dosimetry in the San Juan de Dios Hospital
The event and its discovery
Notes: The information available in his treatment charts was more consistent with good practice

## Slide 10: The discovery of the mistake
On 27 September 1996, after one month working with the new source, the person responsible for dosimetry contacted the physicist at another hospital and asked him to measure the dose rate
The value obtained was 2.02 Gy/min (at the point of maximum dose), while the value that had been used for treatments had been 1.22 Gy/min

## Slide 11: According to the information provided ...
From discussions upon this cross-check, it emerged that during the beam calibration upon source exchange …
…the reading of the irradiation timer had been confused
The value of 0.3 minutes (18 seconds) was taken as 30 seconds, thus resulting in an overestimation of the treatment time by about 66%
      30/18 = 1.66

## Slide 12: The outcome
As a result the doserate was underestimated by the same factor and therefore the dose to patients would have been higher than intended
Subsequent treatment times calculated with the lower dose rate were about 60% longer than required
115 patients were affected and two years after the event, at least 17 patients had died from the overexposure

## Slide 13: Immediate action
The person in charge of dosimetry at San Juan de Dios contacted the Section for Control of Ionizing Radiation of the Ministry of Health of Costa Rica
Ministry immediately ordered unit to cease operation and initiated an investigation
Notes: This section was responsible for monitoring compliance with regulations on radiation protection
It was found that the person in charge of dosimetry had no evidence of an academic degree and had only attended a number of training courses and fellowships

## Slide 14: Use of linear quadratic model
Since fraction dose was higher than normal, the higher fraction size might have an influence on late effects
This effect was estimated for some selected patients
Linear Quadratic (LQ) model was used to calculate the biologically effective dose (BED) and the dose that would produce the same effect, if it were given in 2 Gy fractions
Notes: Doses are given in table form in the following slide

## Slide 15: Some example results for BED and equivalent to 2 Gy/fraction
Notes: <footer>
Linear Quadratic Equation was used to calculate the dose delivered as per the previous lecture

## Slide 16: Error recalled
The value of 0.3 minutes (18 seconds) was taken as 30 seconds, thus resulting in an overestimation of the treatment time by about 66%
Subsequent treatment times calculated with the lower dose rate were about 60% longer than required
Total dose and dose per fraction were increased - increased late effects
Notes: <footer>
Dose increase may result in disproportionate increase in chronic effects. Cannot therefore rely on acute effects for prediction of late effects as this results in an underestimation of the problem
An additional error of 2cm in the optical distance indicator was also identified and they were treating at 78cm rather than 80cm (1.053)

## Slide 17: Potential effects
Cranial and spinal over exposure
  Significant neurological consequences including quadruplegia
Pelvic over exposure
  Severe bloody diarrhoea
Limited lifespan
  Tumour related or further complications
Quality of remaining life and life expectancy may have been significantly reduced
Notes: <footer>
Requires major psychosocial support
May necessitate wearing of diapers which is damaging to self esteem and daily living becomes very difficult
For children receiving CNS irradiation between 20-65Gy in 2Gy/d fractions may on routine clinical investigation be found to perform poorly at school when compared to the peers and may suffer with some level of disfunction of the pituitary gland and hypothalamus
Overexposure – adverse effect may be severe and include lethargy, ataxia, spasticity and progressive dementia, this is potentiated by chemotherapy
Necrosis is serious and irreversible with moderate probability with 40 in 10fr, 50 in 20fr, 60 in 30fr or when fraction size exceeds 3Gy. High probability with 50 in 15fr, 60 in 20fr, or 70 in 30fr
Brain necrosis is manifested by headaches, raised intercranial pressure, seizures, sensory deficits, psychotic changes
Skin changes include telangectasia, fibrosis and chronic ulceration

## Slide 18: Assessment of the impact
Assessment was carried out at 10 and 25 months
 | July 1977 | August 1998
Total patients evaluated | 73 (70 examined) | 53 (51 examined)
Severe/catastrophic effects | 4 | 2
Marked effects / at high risk | 16 | 12
Not severe / low risk | 26 | 22
RT discontinued (underexposure) / at risk of recurrence | 3 | 2
Deceased / radiation related | 7 | 13
Deceased / possibly related | 0 | 4
Deceased / not related | 22 | 35
Deceased / insufficient data | 13 | 9
Notes: <footer>
NB RT discontinued: the effect of stopping RT in these incidences - under rather than over exposure - risk of recurrence. This was an additional unappreciated problem. Treatment had been stopped on discovery of the problem and had not resumed. Some patients were still on active treatment and at least two patients did not resume treatment, one by choice
The side effects observed resulted from overexposure of specific sensitive tissue or a diminution in vascular blood supply. (the most chronic effects of radiation exposure are consequences of an irreversible narrowing of the lumen of arterioles, the thickness of the walls is increased thereby diminishing the size of the lumen and reducing vascular supply) The tissues become thin or atrophic and if severe diminution necrosis may be progressive and may continue years after exposure
In this accident total dose and dose per fraction were increased. May result in a disproportionate increase in chronic effects and you cannot therefore rely on acute effects for prediction of late effects (underestimate)

## Slide 19
No running record of tumour dose or the time for delivery of each fraction
No indication of who treated the patient on a given day
It was not always possible to know which fractions had been used on which machine
No indication of whether charts reviews were performed during the course of therapy
The charts for the other hospital (Hospital Calderón Guardia) were better documented
Difficulties in retrospective dose assessment
Notes: In some instances two times were given, one for each Cobalt unit, but as there was no daily record of the time it was not always possible to know which fraction had been used on which machine

## Slide 20: Causes of the error
Arithmetical mistake
  The value of 0.3 minutes (18 seconds) was taken as 30 seconds, thus resulting in an overestimation of the treatment time by about 66%
  Subsequent treatment times calculated with the lower dose rate were about 60% longer than required
Contributory factors and root causes
  Inadequacy of hospital’s radiation protection programme
  Inadequacy of staffing levels and expertise
  Lack of
      QA programme
      accident prevention measures
      an education and training programme
Notes: As noted in a previous slide the person in charge of dosimetry had no evidence of an academic degree
None of the people in charge at the time of the accident, including the dosimetrist, radiation therapist from the San Juan de Dios and the person who performed the initial calculation and calibration were available for an interview in an effort to reconstruct the charts

## Slide 21: Other small errors that were combined with the main error …
The main error was combined with a number of smaller errors that were present, even before the event. These were:
  Errors in the computer home-made programme to determine the treatment time (about 5%)
  Error of 2 cm with the optical distance indicator - treating at 78 cm rather than 80 cm (1.053)
  Confusion between kerma in air and absorbed dose to water at dose maximum
Notes: This slide summarises some of the additional smaller errors that were present before the incident and that have been referred to previously in this presentation

## Slide 22: Radiotherapy protocols
Expert team noted differences in the radiation therapy practices and protocols for the same disease
Some protocols involved very large fields with the treatment of each field every other day
More than half of the prescribed radiation treatments had fewer than normally accepted number of fractions
Notes: <footer>
These practices undoubtedly aggravated some of the adverse radiation effects

## Slide 23: Clinical consequences
Child affected by overdose to brain and spinal cord
Loss of ability to speak and walk
Notes: <footer>
This and the following slide are photographs of two of the patients who received overdoses as a result of this incident. Minors experienced the highest death rate in the initial stages of the accident. Of the 113 charts reviewed 9 were children

## Slide 24: Clinical consequences
Young woman who became quadriplegic as a result of accidental overexposure to the spinal cord

## Slide 25: Public impact
Loss of confidence in the health care system
  Inaction relative to complaints and problems
  Loss of confidence in some physicians as a result of being ignored or taken too lightly when they complained of severe side effects during treatment
The Minister of Health notified the public very promptly of the accident and was commended for this
Notes: The profile of the patients who were affected showed that they were mostly of modest financial means with a low level of education living principally in rural and suburban areas. The majority (60%) women who did domestic chores and most of the men were farmers. The average number of children per family was 5.
Psychological problems related to a) issue of radiation exposure without somatic effects, b) issues related to having been involved in an accident, c) issues related to the direct effects of radiation on tissues
Patient concerns often related to the unfamiliarity of radiation.
Resolved somewhat by information and education.
People involved in an accident are often thought to have post traumatic stress disorder – with radiation it might be chronic environmental stress disorder
Patients have to live with the knowledge that the potential effects may not be over, or even evident, for years
Patients had formed a self help committee
Patients were rational but nevertheless concerned for their future and the extra risks and possible side effects

## Slide 26: Review of the findings
No records were kept on calibration of radiation beams
No information available on
  equipment used
  the working conditions prevailing
  The dose determination protocols followed
  Results obtained
  Calculation of output/absorbed dose in a computer programme written for calculation of the treatment time
Notes: Given the importance of this incident and the numerous lessons to be learnt the findings are repeated and reinforced on this slide and the following two slides

## Slide 27: Status of charts
No documentation on how calculations of the time had been made following source change
No indication that any review had been made of the treatment time calculation
No running record of
  Tumour dose
  Time for each fraction
  Who treated the patient on any given day
  Portal films acquired - when and by whom
  Chart review during treatment
Notes: <footer>
In some instances two times were given - one for each cobalt unit but since there was no daily record of the time it was not always possible to know which fraction had been used on which machine.

## Slide 28: Failure to …
… perform independent calibration
… notice that treatment times were too long for a new source with higher activity
Notes: There was a failure to perform independent calibration in this hospital but there was also a failure on the part of all involved staff to notice that the treatment times were too long for a new source with higher activity. This should have been seen immediately by the radiation technicians who were working on the unit and also by the radiation oncologist reviewing the patient.

## Slide 29: Facts
Radiation accidents with severe and even fatal consequences do occur in medical facilities
  Human error is the most common cause of radiation accidents
Prior to the accident (external audit had already detected)
  Poor quality record keeping
  Lack of redundancy in procedures
  Inadequate education and training

## Slide 30: Lessons to learn
Carry out an investigation if the results of audit indicate a discrepancy
If there is a high incidence and severity of acute effects it must be investigated

## Slide 31: Lessons to learn
Ensure a high level of training and competence in order to deal with potentially hazardous sources
Specific training should be additional to basic education and not simply attending occasional short courses

## Slide 32: Lessons to learn
Radiotherapy records should be
  Uniform
  Clear
  Consistent
  Complete
Early and reliable information and communication is crucial for good management

## Slide 33: Recommendations – Expert Group to the Government
Radiation protection regulations should be implemented, enforced as soon as possible and kept up to date
QA programmes should be developed and implemented
Education and Training for radiotherapy staff should be improved

## Slide 34: Recommendations – Expert Group to the Government
Record keeping in radiotherapy charts should be improved
External audit - where persisting problems are revealed - another channel of communication to the authorities should be sought
Registry of data on patients should be set up

## Slide 35: Clinical Recommendations
Follow up visits should be structured to allow for
  Optimal individual care
  Scientific evaluation
Baseline data should be compiled according to internationally recognized systems
  RTOG/ESTRO normal tissue effects
  ICRU doses/fractionation

## Slide 36: Reference
IAEA: Accidental Overexposure of Radiotherapy Patients in San José, Costa Rica (1998)
rpop.iaea.org