## Slide 1: Module 2.8: Accelerator interlock failure (Poland)
IAEA Training Course

## Slide 2: Where are we going this time?
Białystok

## Slide 3: Poland - Białystok

## Slide 4: The Neptun 10P Linac
The Białystok
Machine
Built on license from CGR, France by
The Institute of Nuclear Studies, Experimental Establishment for Nuclear
Equipment, Swerk, Poland
1970’s type design
The circuits involved in this accidental exposure are essentially unchanged from the original version
Standing wave type
3 GHz
2 MW pulse magnetron

## Slide 5: What happened?
February 27, 2001
Power failure at the department
Five patients remained to treat that day
Machine was restarted
All machine tests completed without any error indication
x

## Slide 6: What happened?
Analog dose rate indicator fluctuated around 150 MU/min, instead of the selected 300 MU/min
Physicist adjusted the timer to a longer time because of the lower indicated dose rate
He noted a minor beam asymmetry and readjusted for correction
The console of the Neptun 10 P in Bialystok

## Slide 7: What happened?
All 5 remaining patients were treated
  All had 8 MeV electrons
Patients Nos. 3, 4 and 5 soon reported abnormal skin reaction
Patient 5 returned to the radiotherapy department complaining of an itching and a burning sensation
Radiation oncologist also noted erythema which was abnormal
The machine was taken out of clinical use after the last patient

## Slide 8: Action of the physicist
Physicist did measurements
Reading was off scale
Dose rate, without correction for recombination, was
  37 times higher than normal (for 8 MeV electrons)
  17 times higher (for 10 MeV electrons)
  3.5 times higher (for 9 MV photons)
The Neptun 10 P in Bialystok

## Slide 9: Action of the physicist
Physicist noted increased current in filament of electron gun (from 1.20 to 1.46 for 8 MeV)
The accelerator indicated low dose rate
Electronic cabinet

## Slide 10: Vendor came in the next day
Broken fuse
  no power to dosimetry system
Diode broken in interlock chain
  indicates problems in dosimetry system
Low signal from ion chamber
  gun current increased to compensate the low dose rate

## Slide 11: Steps to initiate radiation
Sequence of steps to initiate irradiation includes a test of beam monitoring chambers, but …
… the information about missing power supply can not pass through faulty diode …
… interlock is not informed that monitoring chambers are missing
… and gives green light to the next step in the sequence towards irradiation

## Slide 12: Dose rate vs. gun current
Notes: Normal working condition for the machine is around 1.10 A as filament current, however since the ion chamber signal was very low, the control system assumed no beam or low dose rate and increased the filament current until a sufficient signal came from the chamber. The current increased to approximately 1.45 A which was noted during the verification measurements by the physicist. Measurements by the investigating team have estimated the dose to be somewhere around 70-110 Gy/min in the filament current interval.
There was a limiting level set for the filament current but that was even higher and did not in this case have any influence.

## Slide 13: Estimated patient doses
Notes: The steps to estimate the given doses will be presented.

## Slide 14: Linearity of the monitor chamber
Due to limited equipment
Measurements were done with 25 MU
The linearity of the monitor chamber was studied
Test | Ratio of dose at 150 MU to dose at 25 MU x 6
Normal | 1.00 (Łódź)
Fault 1 /1.48 A) | 1.09 (Białystok)
Fault 2 (1.46 A) | 1.05 (Białystok)
Fault 3 (1.10 A) | 1.09 (Vendor)
Average 1-3 | 1.08
Notes: Linearity problems of the monitor chamber was observed when the polarizing voltage decreased. The graph shows tests done at three locations.
The “normal” is performed at the vendors site and shows a good behaviour (blue line).
The measurements under fault conditions were
performed with the fuse removed and the D29 diode disconnected.
The filament limit was set at 1.48 A (for ‘Fault 1’)
and at 1.46 A (for ‘Fault 2’) in Białystok,
and at 1.10 A (for ‘Fault 3’) in Świerk.
For each measurement series, the doses per MU for the different MU numbers that were pre-set at the console of the accelerator were normalized to the dose per MU at 150 MU.
Thus estimated doses have to be scaled with 1.08

## Slide 15: The saturation in the measuring chamber
1.3 cGy/pulse
ps=1.08
Notes: These measurements using the two-voltage method are verified with theoretical calculations and a correction factor of 1.08 results.

## Slide 16: Reconstruction of fault condition
Measurements made with the equipment in fault condition: without fuse and interlock diode
Filament current at 1.46 A
Made in December 2001
Using three independent methods:
  Ionization chamber
  Alanine
  GAFchromic film

## Slide 17: Measurements in fault condition
Dosimetry method | Dose (Gy) per 150 MU
Plane parallel ionization chamber (Roos) | 127 ± 4
Alanine | 122 ± 4
GAFchromic film | 115 ± 3

## Slide 18: Dose reconstruction from bone samples
Three patients undergoing surgery
Bone samples taken
Dose determined by EPR
Uncertainty: it is not known whether the sample was from the front part or the distal part of the ribs
The dose estimation is done at dmax for both hypotheses
Notes: EPR = Electron Paramagnetic Resonance

## Slide 19: Patient doses
 | Patient 3 | Patient 4 | Patient 5
From EPR |  |  | 
Sample at frontal position | 59 ± 7 Gy | 64 ± 11 Gy | 71 ± 3 Gy
Sample at Distal position | 67 ± 8 Gy | 84 ± 19 Gy | 68 ± 5 Gy
From Ionization chamber | 103 ± 9 Gy | 83 ± 9 Gy | 103 ± 9 Gy
Notes: Biological dosimetry!

## Slide 20: Results on the overexposure
Notes: To follow are several slides showing the clinical consequences of the overdosage for each patient. Details regarding this can be found in the accident report from IAEA.

## Slide 21: Absorbed Doses to the Patients
Patient 1: 50 Gy 60Co +2.5 Gy 8 MeV+ ?
Patient 2: 48 Gy 8 MeV + ?
Patient 3: 25 Gy 8 MeV + ?
Patient 4: 42 Gy 8 MeV + ?
Patient 5: 5 Gy 8 MeV + ?
Patients in order of severity and treatment on day of accident
Notes: An opinion is that only one of these cases have to be shown, e.g. #3 with the CT showing the necrotic area in the thoracic wall; or letting all of this be presented in the “Clinical Aspect” lecture which preferable is given by an physician.

## Slide 22: Patient 1
Dose 50 Gy 60Co + Boost 1x2.5 Gy 8 MeV+ accident
June 4, 2001
Necrosis in area of prior surgical scar
Oct. 2001
Dec. 1, 2001

## Slide 23: Patient 2
Dose 48 Gy 8 MeV + ?
June 4, 2001
White border of lesion
Surgical scar
Oct. 2001
Dec. 1, 2001

## Slide 24: Patient 3
June 4, 2001
October 2001
Dose 25 Gy 8 MeV + ?
Scar, future necrosis
White border
Dec 1, 2001

## Slide 25: Patient 3 - CT of the thoracic wall
Necrotic lesion

## Slide 26: Patient 4
Dose 42 Gy 8 MeV +?
June 4, 2001
Scar, future necrosis

## Slide 27: Patient 5
Dose 5 Gy 8 MeV+ ?
June 4, 2001
October 2001
Dec 1, 2001

## Slide 28: Lessons and recommendations

## Slide 29: Summary
A fault in a fuse of the power supply to the beam monitoring system led to a high dose rate, even though the display indicated a lower value than normal
At the same time, the safety interlock failed
The filament current limitation was set at a high value

## Slide 30: Summary
The probability of double fault was increased because…
… an inoperative interlock could go unnoticed until the second fault appeared
Therefore, the equipment was “ready for the second fault”

## Slide 31: Lessons: Manufacturers
Compliance with IEC safety standards
Review of safety features of existing equipment when a new IEC standard is issued
Explicit recommendations to users on procedures in the case of power cuts (tests to be performed before resuming operation)
Training for maintenance engineers including lessons from accidental exposure

## Slide 32: Lessons: Manufacturers/maintenance
Certification for maintenance engineers should specify restrictions to handle or adjust certain critical parts in the accelerator, depending on the degree of training
Warning notices adjustment of limits to filament current and other safety critical elements
Restricted access to safety critical adjustments be restricted to maintenance engineers certified by the manufacturer

## Slide 33: Lessons: Radiotherapy departments
Immediate check
  upon power supply shut downs or
  any unusual display of dose rate or beam asymmetry or…
Written procedure to ensure that this check is done
If there is a hospital maintenance engineer for the accelerator
  be aware of the limitations, according to certified training by the manufacturer

## Slide 34: Lessons: In short
React and investigate when patients show unusual reactions
QC program must include routines to check accelerator performance after power failure
Equipment should be retrofitted or replaced when technology is out-dated
  This is actually a very complicated process
    who decides and when should it be done

## Slide 35: Reference
IAEA: Accidental Overexposure of Radiotherapy Patients in Białystok (2004)
rpop.iaea.org