## Slide 1: Module 6.3: Source not under control (Turkey and Thailand)
IAEA Training Course
Notes: As these two incidents were very similar with many of the same lessons to learn, the themes have been taken and dealt with in a single lecture

## Slide 2: IAEA
Authorized to
  Establish standards for radiation protection
  Monitor safety of sources of radiation
  Assist in their application
Notes: This slide is to give the role of the International Atomic Energy Agency and its role and authorisation in relation to radiation safety, in this instance, as it applies to active sources and their maintenance, storage and disposal

## Slide 3: BSS
The International Basic Safety Standards for Protection against Ionizing Radiation and for the Safety of Radiation Sources
    Establish the requirements for protection and safety
    States have adequate legal and regulatory infrastructure for effective application of the requirements
    IAEA Safety Standards Series
Notes: <footer>
Requirements and guidance for the establishment of an appropriate infrastructure and other relevant information are issued in the IAEA Safety Standards Series

## Slide 4: Old teletherapy sources
High activity
Regulatory authority must be informed
Requires specialized transport and storage container
Disposal
  could be very costly
  should be part of initial purchase or source replacement contract with manufacturer
  documentation (i.e. source certificate) essential
Notes: This slide is to show the size and structure of the source in the context of the accidents themselves and also to highlight the difficulties in some countries due to the very high cost of disposal in particular if disposal is not part of a replacement agreement

## Slide 5: Disposal of used teletherapy sources
Due to their high activity (even when not anymore fit for use on patients) the disposal of these sources poses a potential risk for severe accidents…
In practice this should not a be major problem if the source is replaced - in this case the manufacturer of the source shall also accept the source to be disposed
Notes: As per the previous slide the major problems arise when disposal of the source is not part of the contract for replacement

## Slide 6: Disposal of used teletherapy sources
A potential for accidents
Disposal routes should already be considered when ordering sources
Accounting for sources includes their disposal
Responsibility for sources only ends after they have been safely disposed and disposal has been documented
Notes: Reinforce that the responsibility for the source remains until disposal has been completed and again the importance of good record keeping

## Slide 7: Background: Istanbul accident
1993 - Specialist company used to exchange sources and maintain equipment recorded as having loaded three spent radiotherapy sources into transportation packages in preparation for returning them to the original supplier in the USA
Notes: At this point the correct procedures seem to have been followed

## Slide 8: Background: Istanbul accident
Turkish Atomic Energy Authority (TAEK) officials
  Checked packages
  Inspected warning signs and transport labels
  Issued permission for transport and export
Packages not exported
  In storage from 1993 - 1998
  TAEK were not informed
Notes: Again the TAEK observed correct procedure, the problem arose due to the action of the company
The packages were stored in Ankara in the first instance and then two of the packages were transported to Istanbul

## Slide 9: Background: Istanbul accident
Feb. 1998 two packages transported to Istanbul
Stored in general purpose warehouse on industrial estate
Warehouse full packages moved to adjoining empty premises
Approx. 9 months later premises sold and new owner sold the two packages as scrap metal
Notes: The new owner was not informed and was therefore unaware of the contents of the two packages

## Slide 10: Istanbul accident: 1998
Notes: This is a map of Istanbul to demonstrate the potential scale of the problem

## Slide 11: Background: Samut Prakarn accident
Company held several teletherapy devices without authorization from the Thailand Office of Atomic Energy for Peace (OAEP)
1999
  Teletherapy heads relocated from a leased warehouse to an unsecured storage location without authorization
Notes: Keeping with the context of themes we are now moving to Thailand to the Samut Prakarn Incident.
This incident is somewhat different at this point as the Thailand Authorities were not informed or aware of the relocation of the teletherapy head which still contained active sources

## Slide 12: Background: Samut Prakarn accident
Late January 2000
  Several individuals obtained access to the location and partially disassembled a teletherapy head
    Took unit home and four people attempted to disassemble it further
Notes: Picture of the teletherapy head with the missing source drawer assembly

## Slide 13: Samut Prakarn accident: 2000
1 - Parking lot
2 - Houses
3 - Junkyard
4 - Hospital
Notes: This slide shows the area involved in the accidental exposure including detail of the houses, depository and hospital that were involved

## Slide 14: Background: Samut Prakarn accident
Source did have a radiation trefoil and warning labels
  Individuals did not recognize the trefoil
  Warning in a language they did not understand
Notes: As will be seen in the following slide on the Istanbul incident the problem of understanding the trefoil signage and the warning was the same

## Slide 15: Background: Istanbul accident
Purchasers
  Did not recognize the trefoil labels which were in place
  Broke open shielded containers and unknowingly exposed themselves and several others in a residential area of Istanbul on 10th December 1998

## Slide 16: Istanbul and Samut Prakarn
Failure to adhere to radiation regulations
Lack of understanding of the meaning of the trefoil symbol
Language difficulties
Notes: This slide makes an initial summary of three of the key messages from this lecture

## Slide 17: Istanbul: chronology
13th December 1998
  10 contaminated people fell ill - 6 began to vomit and sought medical assistance
  Pieces of the containers and at least one unshielded source remained in the residential area for a further two weeks - taken to local scrap yard and remained there for further two weeks
  8th January 1999 examining clinician suspected radiation exposure and notified the national authority
Notes: There was a four week delay between the onset of symptoms and the eventual recognition of the cause. This would not be unusual given the rarity of radiation over exposure seen in a general clinic

## Slide 18: Istanbul: chronology
Examining clinician reported suspected exposure to the national authority
Other victims identified
  18 in total
  11 adults, of whom 10 exhibited clinical signs and symptoms of acute radiation exposure
  7 children

## Slide 19: Samut Prakarn: chronology
1st February 2000
  Two of the individuals took the partially disassembled device to a junk yard in Samut Prakarn so that the component metals could be segregated and sold separately
  During the disassembly the source fell out of its housing unobserved by any of the participants
Notes: This slide again refers to the size of the actual source, and given that this incident occurred in a junk yard with a very large quantity of scrap metal it is not surprising that the source fell out and was unobserved

## Slide 20: Samut Prakarn: chronology
Mid-February
  Several individuals began to feel ill and sought medical assistance
  Physician at the local hospital recognized the signs and symptoms
    Suspected radiation exposure
    Reported suspicions to regulatory authority (OAEP)
Notes: <footer>

## Slide 21: Istanbul and Samut Prakarn
Radiation sickness syndrome recognized by physicians in the local clinics
Authorities notified
Investigation initiated
Notes: This slide again summarises the similarity at this point between the two incidents

## Slide 22: Samut Prakarn: investigation
Mid-February cont.
  OAEP personnel assisted by local health personnel searched for the source
  High level of radiation found in vicinity of junk yard and the area was secured
Notes: <footer>
In the following slides we are moving to the investigations carried out in the two areas
Junk yard secured to prevent further access

## Slide 23: Samut Prakarn: investigation
OAEP officers with public health officers initiated a search for two small pieces of metal, one of which was believed to be the source Unsuccessful
Officers then checked the junkyard
  noted increasing radiation levels in the surrounding area
  Radiation level about 1mSv/h at the entrance of the junkyard
  Identified serious radiological accident and called for assistance
Notes: <footer>
Search was for two small metal pieces that had been given to the other scrap collector. This was done in the belief that one of these small metal pieces was the source.

## Slide 24: Samut Prakarn: investigation
Notes: This slide shows the area where the investigation was centred and also the junk yard which gives a visual impression of the scale of the search that was required

## Slide 25: Samut Prakarn: investigation
OAEP officers called an emergency response team
Command and co-ordination post was established in area of junkyard
Radiation and contamination surveys carried out
  Contamination surveys negative
  Radiation levels at a dose rate up to 10Sv/h
  Evacuation considered unnecessary
Efforts to locate the source continued until 20th February when it was recovered and transported to safe storage
  Maximum recorded individual doses for emergency workers 32mSv
Notes: <footer>
Search was for two small metal pieces that had been given to the other scrap collector. This was done in the belief that one of these small metal pieces was the source.
Emergency crew insisted on wearing lead aprons which were ineffective for activities at this level and also slowed down their progress (however psychologically they felt better)

## Slide 26: Samut Prakarn: investigation
The three missing teletherapy units were located in a parking lot in a suburban area of Bangkok
  Confirmation that the source was from one of the units
  The three units removed to safe storage
Notes: Picture of the three teletherapy heads at the car parking lot.

## Slide 27: Samut Prakarn: investigation
Patients suffered
  Nausea
  Loss of appetite
  Loss of weight
  Skin burns
  Hands swollen and darkened
  Localized hair loss
Notes: This slide indicates the effects of the exposure on the persons involved in trying to dismantle the source

## Slide 28: Samut Prakarn: investigation
Notes: This slide gives the chronology of the persons involved, their relationship to each other and the clinical signs and symptoms and outcomes of the exposure

## Slide 29: Istanbul: investigation
One source was recovered from the scrap yard (activity 3.3TBq (88 Ci))
  Undamaged
  No leakage of radioactive material
3rd source remained in Ankara and later taken into custody
  Fully intact
  Sealed and labeled
  Transferred to the waste treatment section
Notes: The next series of slides related to the investigation of the incident in Istanbul again highlighting the similarities between the two incidents
The source was surrounded by scrap metal which did provide some shielding.
The scrapyard workers were evacuated and the area was completely cordoned off and secured by the national authority and the police
All personnel involved in the source recovery were issued with film badge dosimeters and direct reading electronic dosimeters. The doses were reviewed and recorded at regular intervals

## Slide 30: Istanbul: investigation
Response to advice from the Çekmece Nuclear Research and Training Centre (CNRTC) and a media report
  404 people presented for medical or haematological checks
  18 (including 7 children) admitted to hospital
Notes: <footer>
The 404 people attended following a media report of the incident. Fear of exposure was the motivator
This was the first know major radiological accident in turkey and created much anxiety, with many enquiries from the public concerning their health.
The recovery of the source helped to allay concerns but the psychological impact was high and the family involved suffered anxiety and social isolation from friends and family
Accident created a more general awareness that medical and other applications of radioactive materials can also pose a considerable hazard to the public and requires strict regulations

## Slide 31: Istanbul: investigation
    10 adults exhibited clinical signs and symptoms of acute radiation syndrome
    5 (severe) hospitalized for 45 days
    1 person had signs of radiation induced skin injury on 2 fingers
Notes: This slide shows one of the clinical consequences resulting from very close contact with the actual source within the head, the relationship is clear from the two pictures

## Slide 32: Istanbul: IAEA response
11th January 1999 IAEA Emergency Response Centre alerted to potential accident via routine monitoring of new media reports
  ERC contacted TAEK who confirmed the incident reporting that
    Shield broken into pieces by scrap dealer but source capsule not ruptured
    Two scrap dealers injured and receiving treatment
    Source located and safely stored in a controlled area
Notes: The following slides illustrate the IAEA response to the two incidents

## Slide 33: Istanbul: IAEA response
ERC offered assistance under the Convention on Assistance in the Case of a Nuclear Accident or Radiological Emergency
  Sent medical experts who specialize in diagnosis and treatment of patients exposed to radiation as soon as request received

## Slide 34: Istanbul: IAEA response
12th January 1999
  Official request received from Turkish authority
  Team left for Istanbul
  WHO notified
13th January
  Following examination of 15 patients under investigation in two hospitals IAEA medical team sent preliminary report to ERC
    Treatment satisfactory
    Blood samples taken for cytogenic dosimetry

## Slide 35: Samut Prakarn: IAEA response
21st February 2000: OAEP contacted IAEA
  Described the accident and successful recovery
24th February
  IAEA team of experts requested
26th February
  IAEA team dispatched for 1 week
    2 experts in radiation protection and emergency response management
    3 medical doctors

## Slide 36: Samut Prakarn: IAEA response
  31st March and 26th April 2000
    Additional report providing information on the
      total number of unsecured radiotherapy sources
      Total number of victims and their medical status
      Three subsequent fatalities
      Population screening

## Slide 37: Istanbul: additional missing source
14th January
  Notification of a further missing source
  IAEA field team with monitoring equipment leave for Istanbul
16th January
  Extensive radiation surveys carried out. Checks included
    Building where source stored
    House of initial scrap dealer purchaser
    House of scrap dealer’s father-in-law where damaged source had been taken
    Scrap yard
  No radiation above background levels detected
Notes: These slides return to Istanbul and the additional missing source

## Slide 38: Istanbul: additional missing source
17th January
  Monitoring team travelled 90km to east of Istanbul where smelting had taken place
  Extensive radiation surveys carried out but no activity above background levels was detected
Notes: The train of events was retraced and uncovered that the containers had remained in the yard until 27th/28th December when they were loaded by had and shovel onto a truck and taken to a larger scrap metal yard in the area. The owners of this yard were away and returned on 4th January by which time most of the scrap metal had been taken to a large smelting factory and had probably been smelted on 3rd January.

## Slide 39: Istanbul: IAEA team recommendations
Prepare press release to inform public
Provide police and fire brigade with pictures and drawings of the stainless steel canister
Alert medical personnel and hospitals to report any suspicious cases to TAEK
Continue police investigations
Update inventory of radiation sources and equipment in Turkey
Notes: <footer>
The following slides outline the IAEA Team findings and recommendations
Source sold illegally to other unknown person in Turkey or abroad

## Slide 40: Istanbul: IAEA team recommendations
To inspect storage condition of radioactive sources and substances in Turkey
The necessity of informing States which may be affected
Informing the IAEA of missing sources
To continue with radiation survey of all possible sites for the lost source
Install radiation monitor immediately at the metal production factory
Notes: <footer>
Colakoglu metal production factory where smelting took place

## Slide 41: Istanbul: IAEA team recommendations
Interview factory engineer responsible for the nucleonic level gauges
Check factory records
  Sales over previous 9 months
  Purchasers
  Monitoring construction or structures for which metal was intended
  Possibility of sale or exchange with other factories

## Slide 42: Istanbul: IAEA team recommendations
Estimate possible activity concentrations in steel produce
Survey potential land dumps and municipal general dumps
20th January
  Further IAEA expert sent to Istanbul to complement work already carried out
26th January
  Neighbouring states notified of missing source
Notes: <footer>
Through data on size of furnaces and the approx amount of recycled steel produce

## Slide 43: Samut Prakarn: IAEA findings
Difficulties in disposing of disused radiation sources
Lack of communication / notification of authorities on disposal
Failure to fully implement national regulations
Regulations and license requirements unclear
Delay in diagnosis of radiation sickness syndrome

## Slide 44: References
IAEA: The Radiological Accident in Istanbul (1988)
IAEA: The Radiological Accident in Samut Prakarn (1988)
rpop.iaea.org
rpop.iaea.org