## Slide 1: Module 2.2: Erroneous use of TPS (UK)
IAEA Training Course

## Slide 2
Background
Until 1982, a hospital relied on manual calculations for the correct dose to be delivered to the tumour
Treatments were generally performed at standard SSD (100 cm)
SSD = 100 cm

## Slide 3
Background
Isocentric treatments were rarely given in the hospital, because calculations were cumbersome
Isocentric

## Slide 4
Background
Some non-standard SSD treatments were performed. SSD-correction was then applied.
SSD = 120 cm
SSD-correction!

## Slide 5
Isocentric treatment
Machine used principally for SSD treatments would have been calibrated at 100 cm SSD + dmax
For isocentric treatment, patients generally positioned with center of PTV at machine isocentre
isocenter
Calibration in water phantom
Standard SSD
isocenter
Calibration in water phantom

## Slide 6
Isocentric treatment
Using a different SSD, means a change in output factor compared to the standard calibration (and a change in depth dose which can often be ignored)
The change in output factor depends on inverse square law
isocenter
Calibration in water phantom
Standard SSD
Notes: The patient is closer to the source thus the dose rate increases.

## Slide 7
SSD = 90 cm, E = 6 MV
Example:
((100+dmax) / (90+dmax))2
(101.5 / 91.5)2 = 1.23
(Indicating that the dose rate at the shorter distance is 23% greater than at 100 cm SSD)
Calculation procedure
A non-written procedure was in effect for treatments at non-standard SSD (including the few isocentric treatments). Technologists calculated a correction factor based on the actual SSD used.
Notes: This highlights the need for written procedures to be in place in all departments

## Slide 8
TPS installation 1982
A computerized treatment planning system was acquired in 1981, and after some preliminary testing brought into clinical use in autumn of 1982
Partly because TPS simplified the calculation procedures, the hospital began treating with isocentric techniques more frequently

## Slide 9
First isocentric treatment plan from TPS
When the first isocentric TPS plan was ready and presented to the planning technologists, the following happened:
  It was assumed by the technologists that correction factors for non-standard SSD should be applied
  Hospital physicists approved this procedure
Notes: This highlights the importance of good communication within departments and the need for all staff to be involved and aware when new equipment and procedures are introduced.

## Slide 10
First isocentric treatment plan from TPS
It was not recognized that the TPS already correctly applied an inverse-square correction for isocentric treatments!
Notes: Failure to understand the functionality of the new system and to carry out full commissioning at the time of installation

## Slide 11
The technologists continued to apply the distance correction factor to all subsequent calculations
Consequently, distance correction factor was applied twice for all patients treated isocentrically, or at non-standard SSD
This error caused patients to receive doses lower than prescribed
Subsequent isocentric treatment plans
Notes: All isocentric plans from this time 1982 – until discovery in 1991 were calculated in this way
Given that lower doses were delivered it was possibly more difficult to detect than an overdose as reactions would be less severe

## Slide 12
Discovery of error
In 1991 a new computer planning system was installed and a discrepancy was discovered between the new plans and those from the previous system
Further investigation revealed that the original TPS already contained within it the correction for calculations at non-standard SSD.
  Systematically re-applying the correction factor resulted in underdosage
Notes: Underdosage - between 5% and 35% of the prescribed dose actually delivered

## Slide 13
A formal investigation was initiated
The incorrect procedures were in place until 1991, or for approximately nine years
During the 9-year period, 6% of patients treated in the department were treated with isocentric technique; for many of these patients it formed only part of their treatment
Investigation of error
Notes: The impact was lower as it only applied to the patients receiving isocentric treatments
The patient group most affected were the patients treated radically with radiation only
It was less marked for patients who were also having surgery and for patients with metastatic disease

## Slide 14
All patients receiving isocentric treatment (performed on two linear accelerators) between Autumn 1982 and December 1991 were identified
Evaluation by Ash and Bates showed that of 1045 patients whose calculations were affected by the incorrect procedures, 492 developed local recurrences that could be attributed to the error
Underdose varied between 5 and 35%
Evaluation of error

## Slide 15
Dose reduction distribution for patients
Notes: This is a graphical representation of the dose reduction with the majority of patients underdosed by between 11% and 30%

## Slide 16
Patient identification
Data stored on floppy discs had become unreadable due to age
Instead: systematic examination of log books for each of the two linear accelerators was necessary
  Log book records SSD for each treatment
  Patients with SSD < 100 cm were identified and doubly checked by referring to their treatment plan
Notes: Unlike the previous incident discussed accurate data in this case was not readily available.
To identify the patients the log record books for the two linear accelerators had to be used. In these log books FSD was recorded and therefore all FSD below 100cms were identified and the patient charts located) this search identified 989 patients, 445 of whom were alive and 544 dead.
Public disclosure was necessary to identify any additional patients who may have been affected. This resulted in 4930 telephone calls from the public with 470 of these from patients who had received isocentric treatment. These were then checked against the 989 already identified.
All patients attending the hospital for follow up were also checked and the two actions identified a further 56 patients
The total number of patients therefore was 1045

## Slide 17
Data reviewed
Patient identification
Diagnosis
Stage
Grade
Treatment details
  Prescribed dose
  Shortfall in dose actually delivered
Outcome
  Survival
  Patterns of recurrence
Notes: A detailed analysis then took place to try to establish outcomes for the affected patients. This was an independent review of the case notes with data entry onto a proforma sheet. The data entered is listed on the slide
A clinical judgement was then made on the likely contribution of the underdosage to the final outcome. Such an analysis required detailed information on the individual patients and their cancer status at the time of treatment. In these cases the patient case notes were very good and it was possible to assess the histological type of the tumour and the stage and grade.

## Slide 18
Clinical impact
Based on the relationship between radiation dose and symptom control
Difficult to asses the impact given the complexity of the factors affecting tumour growth, development and response to treatment
Post mortem data was not available and information on death certificates may be unreliable
Notes: Theoretically there is a correlation between radiation dose and symptom control but this is difficult to assess categorically due to the complexity of individual factors in each patient.
An estimate of the likelihood of cure related to dose was based on data from other centres on similar patient groups. Interestingly this resulted in a range of survival probability with a variation across centres.
Of the patient population evaluated a number of patients had already died and some patients had been treated too recently to fully assess the effect of treatment on outcome.
There was some reliance on death certificate data and this is often unreliable (patients may die from pneumonia for instance as a result of their underlying malignancy but on the death certificate it will be recorded as pneumonia and not cancer)

## Slide 19
Clinical data
Site | Patient profile | Treatment intent | Dose reduction | Outcome
Bladder | 242 | Radical 236 Palliative 3 Post operative 3 | 21-30% in 204 pts | 39 no effects 46 uncertain 150 adverse affect possible
Cervix | 162 | Radical 160 Palliative 5 Post operative 30 RT alone 127 | 21-30% in 134 pts | 43 no effects 46 not assessable 70 adverse affect possible
Endometrium | 104 | Post operative 88 RT alone 13 | 21-30% in 83 pts | 42 no effects 46 uncertain 13 adverse effect possible
Lung | 206 | Radical 190 Palliative 16 | 11-20%in 103 pts | 79 no effects 15 uncertain 100 adverse effects possible
Oesophagus | 134 | Radical 131 Palliative 3 | 11-20% in 95 pts | 34 no effects 10 uncertain 82 adverse effects possible
Rectum | 75 | Pre or post operative 25 Recurrence 37 Palliative 9 Unknown 4 | 21-30% in 48 pts | 25 no effect 41 adverse effects possible
Prostate | 47 | Radical 42 Palliative 5 | 21-30% in 39 pts | 20 no effect 9 uncertain 23 adverse effects possible
Notes: This chart details the 1045 patients evaluated and identifies the patients where an adverse outcome was considered a possibility
A local recurrence within three years from treatment was taken as the normal critieria for a relationship with underdosage and for completeness the total number of patients with recurrence were included in the analysis
Some patients had gone on to have salvage surgery which reduced the impact but may not have been necessary under normal conditions
996 patients were treated with radical intent, 49 palliative. There were no children, breast, head and neck or brain cases involved

## Slide 20
Actions advised
Patients dead
  Information and counselling for family
Patients alive
  Follow up with short intervals
  ? Further radiation to make up for missing dose
    Radiation completed 1 – 2 months before discovery: YES
    Radiation completed > 3 months before discovery: NO
  Radical surgery
Notes: 640 patients were dead and 393 alive at the time of the analysis. The outcome was unknown in 12 patients
A significant degree of distress was caused to both surviving patients and their relatives and also to the relatives of the patients who had subsequently died. It was necessary to provide psychological support to families of the patients who had died.
For surviving patients the important issue was to try to compensate in some way for the underdosage and the range of actions listed in the slide were put into place

## Slide 21
Problems highlighted
Lack of communication between the professional groups involved
Failure to fully evaluate the new TPS
Lack of education
Failure to implement correct policies and procedures
Lack of independent checks within the system
Patients often followed up by non-radiotherapists (e.g. urologist)
Notes: This case again highlighted the need for good communication, written procedures and protocols and independent checks.
It also raised the issue of patient follow up which is often not carried out by personnel familiar with the side effects associated with radiotherapy. There was also a failure to report back to the oncologist after follow up review

## Slide 22
Lessons: Radiotherapy Department
Ensure that staff are properly trained in the operation of the equipment
Ensure that staff understand the operating procedures
Include in the Quality Assurance Programme:
  Procedures to perform complete commissioning of treatment planning equipment before first use
  Procedures for independent checking of patient treatment time calculations

## Slide 23
Lessons: Radiotherapy Department
Importance of reliable and comprehensive databases
Need for follow up of patients by clinicians with a background in radiotherapy
Notes: Unlike the previous case, as stated at the beginning, the data was difficult to access. A large number of patients suffered unnecessary distress during the attempt to identify all affected patients (4930 phoned the hospital following the public announcement)
The patients charts however were detailed for the time and this facilitated an analysis of possible outcome
It highlighted the need for follow by personnel with a background in radiotherapy but this is very difficult to establish in some countries given the staff shortages and heavy workloads

## Slide 24
Lessons: Radiotherapy Department
Underdose is difficult to asses as it does not produce recognizable symptoms
Audit of outcome
  Overall survival
  Disease free survival
  Local recurrence rate (related to stage and grade of the cancer)
  Publication of results at regular intervals
Notes: This was further compounded by the fact that underdosage does not produce recognisable symptoms and in many of these cases the underdosage was such that the patients did experience some side effects
Audit of outcome is more likely to be carried out in clinical trials settings
Publication of results at regular intervals should be encouraged as this will identify major deviations from national norms but this requires comprehensive clinical information systems and trained staff

## Slide 25
Reference
Ash D, Bates T. Report on the clinical effects of inadvertent radiation underdosage in 1045 patients. Clin Oncol 6: 214-225 (1994)