## Slide 1: Module 2.5: Incorrect repair of accelerator (Spain)
IAEA Training Course

## Slide 2: Hospital Clinico – Zaragoza - Spain

## Slide 3: Events: an overview
5th December 1990
  no electron beam on linear accelerator
  noted in the log containing data regarding the daily treated patients as: “11:30; breakdown”
A technician was at place from General Electric-CGR
  Maintained a 60Co unit at the clinic
  The clinic had a maintenance contract with GE/CGR
  The technician had a first look and decided to postpone the work until the next workday
Notes: The problem was communicated to a technician from General Electric-CGR
He was maintaining a Co60 unit in an adjacent room at the clinic
The clinic had a maintenance contract with GE/CGR
The technician had a first look and decided to postpone the work until the next workday

## Slide 4: Events: an overview
6th December 1990 – Holiday
A repair was carried out by the technician the following day
  the beam was recovered but …
  …, an instrument on the control panel always indicated the maximum electron energy (36 MeV), regardless of the selected electron energy value 7, 10, 13 MeV etc
Treatments resumed Monday the 10th December

## Slide 5: A “faulty display”
The technologists observed the discrepancy between the energy selected and the one indicated on the instrument on the control panel
The interpretation was
  (the needle) “must have got stuck at 36 MeV” but
  the energy must be as indicated on the energy selection keyboard
Notes: <footer>
The speedometer in this slide is not the energy display from the console of the linac, however, a similar type of meter was present to indicate the energy of the beam. (This meter here is from a Chevrolet in the 50-ies!)

## Slide 6: Events: an overview
20th December
  the Physics and Radiation Protection Dept. is informed about the incorrect energy display
The linac is immediately taken out of service
  observe - after 10 days of treatment
Physicians start to correlate the low tolerances and the reactions among patients with the event
Notes: One may note that it takes 10 days until the physics department is informed about the incorrect display on the linac.

## Slide 7: Events: an overview
At this point, no information was given to the maintenance service of the hospital about
  The original breakdown of the linac
  The repair by the technician
This information was given a month later on the 20th Jan 1991
Notes: Still notable that the repair has not been communicated to the physics department. The organization maybe questioned as is well outlined in the report from the Spanish Society of Medical Physics.

## Slide 8: Events: an overview
*CSN - Consejo de Seguridad Nuclear
21st December
  Dosimetry checks reveal the energy is 36 MeV! regardless of selection on the control desk…
The company is informed and sends a technician to investigate and repair
Investigation by CSN* on the 5th Jan. shows:
  7 MeV - Dose increase 7 times
  10 MeV - Dose increase 5 times
  13 MeV - Dose increase 3 times

## Slide 9: Consequences: an overview
During the 10 days
  27 patients were treated using electrons with the faulty equipment
Of the 27 patients
  15 died as a consequence of the overexposure
    Most of them within 1 year
    Radiation injuries of the lung and spinal cord
  Two more died with radiation as a major contributor

## Slide 10
 |  |  | Clinical findings or Cause of death | Death | Radiation
MV | 33 | F | Radiation induced respiratory insufficency | 1991-05-20 | Yes
BC | 69 | F | Rupture of esophagus due to overexposure | 1991-05-08 | Yes
PS | 45 | F | Myelitis, paraplegic, esophageal stenosis | - | Yes
DR | 59 | F | Pneumonitos, hepatitis due to overexposure | 1991-03-26 | Yes
JC | 60 | M | Hypovolemic shoch due to radiation induced hemorrhage in neck | 1991-09-14 | Yes
FT | 68 | M | Myelopathy due to radiation | 1991-04-15 | Yes
MP | 55 | M | Myelopathy, lung metastases, respiratory insufficiency possibly due to radiation | 1991-03-16 | Yes
IL | 65 | M | Myelopathy postradiation | 1991-12-25 | Yes
JV | 67 | M | Left thigh and groin fibrosis |  | 
AS | 67 | M | Ulcerated hypopharynx, cervical myelitis, radiation burn of neck |  | 
JG | 60 | F | Respiratory insufficiency due to overexposure | 1991-09-07 | Yes
AG | 60 | F | Respiratory insufficiency due to overexposure | 1991-07-28 | Yes
BG | 50 | F | Healed skin burns of anterior chest |  | 
CM | 51 | F | Respiratory insufficiency due to overexposure | 1991-03-09 | 
AR | 71 | F | Skin burns, esophagitis, femoral vein thrombosis | 1992-04-08 | Probably not
IG | 68 | F | Paraneoplastic syndrome, metastases | 1991-11-22 | No
SA | 45 | ? | Inguinal skin burns |  | 
FS | 59 | F | Pneumonitis and myelopathy | 1991-08-29 | Yes
JS | 42 | M | Skin burns shoulder, fibrosis, necrosis |  | 
TR | 87 | F | Respiratory and renal insufficiency and encephalopathy due to overexposure | 1991-07-12 | Yes
BF | 39 | F | Respiratory fibrosis and metastases | 1992-05-20 | Yes
NC | 72 | F | Skin burns chest, pleural and pericardial effusion |  | 
PS | 42 | F | Respiratory insufficiency due to overexposure | 1991-02-21 | Yes
LS | 72 | F | Generalized metastases | 1991-01-09 | No
JG | 80 | F | Generalized cancer | 1991-01-08 | No
JS | 56 | M | Myelopathy due to overexposure | 1991-02-16 | Yes
SM | 53 | M | Myelopathy due to overexposure | 1991-02-17 | Yes
'From: Accidents in Radiation Therapy, FA Mettler Jr, P Ortiz-Lopez in 'Medical management of radiation accidents, Ed. IA Gusev, AK Guskova, FA Mettler. 'Published by CRC. ISBN 0-8493-7004-3
Notes: This slide is here only to be complete. It also shows how well the follow-up was performed. All the patients and the individual consequences are shown. One can high light certain finding etc…

## Slide 11: Technical and Physical Description of the Event
According to a report from the Spanish Society of Medical Physics
Notes: Thus lets go into the technical and physical description of the accident

## Slide 12: Specifications of the accelerator
Electrons
  7, 10, 13, 16, 19, 22, 25, 32, 40 MeV
Photons
  25 MV
Traveling-wave guide
Bending magnet system - slalom type
No flattening filter
  Beam scanned (up to 36 x 36 cm2)
Notes: <footer>
Specifications of the accelerator, relevant to the accident

## Slide 13: The Sagittaire accelerator
Travelling wave guide
Gantry and treatment head
Images courtesy of Rune Hafslund
Notes: Pictures from the installation in Bergen, Norway.

## Slide 14: The electron path
The path is controlled by electromagnetic field, bending magnet
Higher current needed when electron energy increases
Only one current is correct for a single electron energy (the deflection current)
37°
127°
e-
e-
37°
37°
Notes: <footer>
The electron path through the gantry is displayed. The acceleration is performed in the several metres long wave guide (appr 8 m) and when the electrons enters the gantry they are bent through several electrically controlled bending magnets. For each electron energy a specific current, the deflection current, is used for each magnet to get the electron through the system.

## Slide 15: The electron path
Deflection current through bending magnet coils
  Is too high
  Curvature radius is too short
  Electrons are “lost”
No radiation beam
Correct path
Electrons lost the path
Notes: <footer>

## Slide 16: The electron path
Deflection current is too low
  Radius of curvature is too large
  Electrons depart from the correct path
No radiation beam
Electrons lost the path
Correct path
Notes: <footer>

## Slide 17: The electron path
Correct deflection current through the coil of bending magnet
That matches the electron energy
Electrons will find their path
Then we have a BEAM
Electrons lost the path
Electrons lost the path
Correct path
2
1
Notes: <footer>
When the current matches the energy we will get the correct path. If the current is too low compared to what is required for a certain energy the electrons will not be bent sufficient (1), On the other hand if the current is too high they will be bent to much (2).
But if the energy is too high compared to the current we will also follow path (1) and if the energy is too low path (2) will be the result.

## Slide 18: Equipment fault
Always maximum deflection current
7 MeV
10 MeV
13 MeV
… MeV
Transistor short-circuited,
Notes: <footer>
The switching of energy selects different resistors which then controls the base of the transistor controlling the deflecting current to the bending magnet. Unfortunately the transistor wa short-circuited and independent on the voltage at the base, full current was always fed to the magnet system.

## Slide 19: Equipment defect: maximum deflection current
No electron beam possible (except for maximum energy)
Correct path: only possible with maximum energy (MeV)
Electrons lost the path
7 MeV
10 MeV
13 MeV
Notes: <footer>
The consequence of the broken transistor was that the deflecting current went to maximum, thus a beam could only be found if the energy was increased.

## Slide 20: During the repair
Energy was adjusted until beam was found
  This was done for all energies
Since running at maximum deflection current
  => ~36 MeV for all electron beams
Instead of finding the defective (short-circuited) transistor and restoring the correct deflection current in the bending magnet
To do this adjustment
  energy selection had to be switched to “manual mode”
By doing so, the energy selection from the control panel was partly disabled
Notes: <footer>

## Slide 21: Electron energy
The instrument indicated always 36 MeV
… regardless of whether the selected energy was 7, 10, 13, 16 … MeV
The keyboard for the energy selection disabled
Notes: <footer>

## Slide 22: Field size
A homogenous field is achieved by scanning the electron beam
The current of the scanning magnet has to match the selected electron energy
Notes: <footer>

## Slide 23: Field homogeneity
As the electron energy was at the maximum, the deflection in the scanning magnets was too small and the field was concentrated in the centre
This increased the energy fluence and therefore the dose
Notes: <footer>

## Slide 24: Simulation of fault – homogeneity of beam
Symmetry error could reach 30 %
  Due to symmetry I/L inhibited?
    Or a failure/fault
  Due to symmetry I/L has inadequate tolerances?
Only partial irradiation of monitor chamber
  Worst case only ~45 % was covered => 5x error in dose
Increased gun current to compensate low dose rate
Notes: I/L = Inter Lock
There are some ambiguities in the description of the event regarding the symmetry interlock of the machine. Was it overridden by the technician or not? Was it broken? However, a strongly un-symmetric beam could be delivered! Thus only parts of the monitor chamber was irradiated and according to the investigation this could lead to errors in output (dose per m.u.) up to 5 times higher dose.

## Slide 25
The
  higher fluence of particles combined with
  higher energy of particles
…, caused
  higher absorbed dose
  deeper penetration in the human body
Energy (MeV) E0 | 7 | 10 | 13 | … |  | 35
R50 (cm) | 3.0 | 3.9 | 5.6 | … |  | 15.0
Assuming displayed energy is E0 thus E0=2.33 . R50
Result of higher electron energy and fluence
Notes: <footer>

## Slide 26: Dose excess as function of electron energy
For 7 MeV, the absorbed dose was about 9 times the intended
This increase was smaller for higher energies
It became nearly unity (when the selected energy coincided with the actual energy)
(Based on measurements after the accident)
Actual energy

## Slide 27: Event in summary
An incorrect repair was made
  A fault was corrected by an erroneous adjustment
  The origin of the fault was not investigated
A beam with higher penetration and much higher dose was produced
  ~35 MeV, 9-1x higher dose
The energy indicator was showing that the energy was incorrect
This indication was not analyzed until 10 days of treatments, involving 27 patients
No report to physics about the fault, repair, etc.
Notes: <footer>

## Slide 28: Causes of event
Short circuit in the power supply/regulator of the bending magnet coils
The staff at the linac failed to follow the hospital’s regulations
  Document the breakdown
  Report to the right person
  Requested help instead directly from an un-qualified person
Notes: The unqualified person is the technician who did not have any training on the accelerator in question.

## Slide 29: Causes of event
The technician concentrated on getting a beam (whatever beam) instead of identifying the cause of the fault
  Bypassed some safety checks
Treatments continued without any reporting for 10 days
  Even when the energy meter indicated the wrong energy
Failure to interpret the complaints from the over-irradiated patients as a warning

## Slide 30: Lessons: Radiotherapy Department
Include in the Quality Assurance Programme
  Formal procedures for
    returning medical equipment after maintenance,
    making it mandatory to report it to the Physics group, before resuming treatment with patients
  Consideration of the need to verify the radiation beam by the Physics group, when the repair might have affected beam parameters
  Procedure to perform a full review or investigation when unusual displays or behavior of the radiotherapy equipment occurs
Notes: <footer>

## Slide 31: Lessons: Radiotherapy Department
In case of unusual reactions in one patient notified by a technologist or directly by the patient, the radiation oncologist should immediately request the medical physicist to perform a verification to detect a possible error in any of the treatment steps
Unusual reactions in more than one patient should lead to a request to the medical physicist to immediately verify the dosimetry of the treatment unit
Notes: <footer>

## Slide 32
Aftermath
A GE technician was found guilty of criminal negligence in a Spanish court for his role in what experts are calling the world’s worst radiation therapy accident, in which 27 patients allegedly received overdoses from a malfunctioning radiation machine at a hospital in Zaragoza, Spain during a 10-day period in December 1990.
A Zaragoza judge handed down the decision in April, determining that the overdoses resulted in 20 deaths and seven serious injuries.
According to GE, the court found both the company’s service technician, and GE-CGR España civilly liable for the $3.7 million award to the accident victims. Although the technician was found guilty of criminal negligence, GE-CGR España was not the subject of any criminal charges.

## Slide 33: References
The Accident of the Linear Accelerator in the “Hospital Clínico de Zaragoza”, Spanish Society of Medical Physics. (1991)
El Periódico. Account of the court proceedings and the verdict for the case of the accelerator accident in Zaragoza (Spain) reproduced in the newspaper in April 1993 (in Spanish)
Notes: <footer>