## Slide 1
Module 5.2: Reporting
and investigating

## Slide 2
Background information

## Slide 3
Background
Targeting PTV while avoiding OR
Notes: The first few slides are “back-to-basics” – the intended treatment targets the Planning Target Volume (PTV in red) while avoiding, as much as possible, the Organs at Risk (OR in green). This set of slides is shown in order to illustrate intended treatment versus not intended treatment.

## Slide 4
Background
Targeting PTV while avoiding OR
Notes: A “volume of crossfire” is built up with irradiation beams from various incidence directions.

## Slide 5
Background
Targeting PTV while avoiding OR

## Slide 6
Background
Targeting PTV while avoiding OR

## Slide 7
Background
Targeting PTV while avoiding OR

## Slide 8
Background
Targeting PTV while avoiding OR
Notes: The treated volume, given an absorbed dose high enough to achieve the intended purpose of the treatment, is inside the white line.

## Slide 9
Intended treatment (2D-view)
Background
Notes: We end up with the intended treatment – the PTV within the treated volume.

## Slide 10
Intended treatment (3D-view)
Background
Notes: This is the same plan with more information in a 3D view.

## Slide 11
Intended size and location of treated volume
Background
Notes: This is the intended size and location of the treated volume in relation to the PTV and OR.

## Slide 12
Unintended size and location of treated volume
Background
A mistake in the calculations of the treatment might shrink the treated volume unintentionally.
Insufficient TCP
Notes: For various reasons (for example a mistake in the calculations) we can end up with an unintended size of the treated volume – resulting in e.g. insufficient Tumour Control Probability.

## Slide 13
Unintended size and location of treated volume
Background
A mistake in the transfer of calculated parameters might shift the treated volume to an unintended position.
Insufficient TCP, unacceptable NTCP
Notes: ..another example where the location of the treated volume has shifted unintentionally.

## Slide 14
Unintended size and location of treated volume
Background
There are examples of simple mistakes causing massively different absorbed dose than intended.
Unacceptable NTCP
Notes: An example of a simple mistake causing massively different absorbed dose than intended is the event in Costa Rica.

## Slide 15
Medical errors are a major problem (Institute of Medicine (USA), 1999):
As many as 44,000 to 98,000 people die in hospitals (USA) each year as result of medical errors
This compares to 43,458 for motor vehicle accidents or 42,297 for breast cancer
Comment: It is likely that medical errors occur with at least the same frequency in other countries
Background

## Slide 16
Medical errors are a costly problem (Institute of Medicine (USA) 1999):
There are estimates that medical errors cost the USA around US$ 37,600,000,000 each year
About half of this cost is associated with preventable errors
$
Background

## Slide 17
Background
With hindsight, it is easy to see a disaster waiting to happen. We need to develop the capability to achieve the much more difficult - to spot one coming
DoH An organization with a Memory

## Slide 18
When addressing this: what we need to know:
What can go wrong in the process? (Hazard identification)
    Systematic review of inherent hazards in system
    Many methods can be used, e.g. foresight and review of retrospective data (reported incidents)
How likely is this to happen? (Frequency analysis)
    Determination of frequency of these events
    Retrospective data (reported incidents) – “near misses” have special role since data is more often captured at this stage
What are the consequences? (Consequence analysis)
    Estimation of impact if the event occurs
    Consequence models are required – in radiotherapy e.g. undesired outcome of treatment
The combination of frequency and consequence tells us the risk
Background
Notes: It is of benefit to know the hazards and their characteristics (frequency and consequence) in order to address them properly. When doing the hazard identification, one can use historic data (e.g. incidents that have previously been reported) or “think ahead” of what might happen. When doing the frequency analysis, one should consider that “near misses” (here taken to mean an error that was found before it had any consequence for a patient) might be found with greater regularity than actual errors that affect a patient’s treatment, and would therefore capture valuable data on likelihood of specific errors happening. When doing consequence analysis, the probability models of TCP / NTCP for change in dose could be used – this is however quite a complex task to do.

## Slide 19
When addressing this: what we need to do:
To improve the safety in radiotherapy, we should aim to:
Minimize occurrence of errors (but some errors will still occur)
Find errors before they are causing harm (but some errors will still be present at treatment where they might cause harm)
Minimize the harm caused
Background
Notes: We don’t only need to know about the risks in radiotherapy (previous slide), we also need to do something (this slide).

## Slide 20
When addressing this: what we need to do:
We aim to minimise the risk through multilayered prevention
Background
Initiating events
Accidental exposures
Notes: Participants can be reminded about the previous lecture 5.1 on prevention.

## Slide 21
These layers should encompass:
Actions where potential deviations from intended dose and geometry can be found before the first irradiation fraction of the patient (e.g. chart-checking)
Background
Notes: Different types of “layers of safety” are presented over the next few slides. Incident reporting is suggested to be one of these types of safety-layers.

## Slide 22
These layers should encompass:
Actions where deviations can be found during or after the treatment course (e.g. in-vivo dosimetry)
Background

## Slide 23
These layers should encompass:
Application of safety technology (e.g. integrated radiotherapy networking)
Background

## Slide 24
These layers should encompass:
Actions where contributing factors such as staffing-levels and structure, training and communication are addressed (e.g. monitoring of workload)
Background

## Slide 25
These layers should encompass:
Application of safety procedures (e.g. incident reporting systems)
ROSIS
Background
Notes: Having suggested that an incident reporting system is just one type of safety layer, the module will now focus explicitly on incident reporting, and the characteristics thereof.

## Slide 26
Incident reporting systems:
mandatory or voluntary systems

## Slide 27
Incident reporting systems
Different types of incident reporting systems:
Mandatory reporting systems:
Reporting of certain events is required (e.g. reporting to regulatory authorities of events above certain magnitude)
Voluntary reporting systems:
Reporting is encouraged (e.g. reporting to professional body)
Notes: Two different types of incident reporting systems are the mandatory and the voluntary systems. Examples of these will be given later in the module.

## Slide 28
Mandatory reporting systems
Mandatory reporting (to regulatory authorities) should …
… ensure providers of medical care are held accountable to the public
… focus on serious errors resulting in injury or death
… require reporting of information in a standardized format to a national database
Notes: Focussing first on the mandatory reporting systems, the characteristics of these systems are given in the following slides.

## Slide 29
Mandatory reporting systems
Three main purposes:
Provide public with a guaranteed level of protection by assuring that most-serious errors are reported and investigated, and action is taken
Provide an incentive to hospitals to improve patient safety
Require hospitals to invest in patient safety, helping to assure that hospitals offer comparable care

## Slide 30
Mandatory reporting systems
Filing of a report should not trigger a release of information
Reporting should trigger an investigation
Release of information should occur only after incident has been investigated thoroughly
Information released should be accurate
Employees should feel confident that response to reporting of significant error will be reasonable and justified

## Slide 31
Mandatory reporting systems
Radiotherapy: A mix of radiation and medicine
Legislation and regulations concerning reporting of incidents in radiotherapy can be covered in relation to radiation protection and / or health
In some countries, radiation protection legislation makes it mandatory to report radiotherapy incidents to a higher authority
In some countries, health legislation makes it mandatory to report radiotherapy incidents to a higher authority
Some countries stipulate that local recording of incidents is mandatory. Potential incidents are covered in some countries
Notes: From the point of view of legislation and regulations, radiotherapy can be considered to contain elements of both radiation and medicine.

## Slide 32
Voluntary reporting systems
Voluntary reporting systems should …
… encourage hospitals to focus on improvement of safety environment
… focus on errors that result in little or no harm to patients
Notes: Focussing now on the voluntary reporting systems, the characteristics of these systems are given in the following slides.

## Slide 33
Voluntary reporting systems
Emphasis should be on improving patient safety
Mechanisms should ensure that information and lessons learned can be shared effectively
Reporting system should allow analysis to select most effective means for improving safety
Focus should be on accurate and complete reporting of information to facilitate speedy investigation and action

## Slide 34
Voluntary reporting systems
Voluntary reporting systems should be protected from the legal system
Mechanism should allow handling in confidence
Mechanisms should allow for anonymous reporting of errors or circumstances that could lead to errors
Staff should not fear punishment
Signalling likelihood of retribution or punishment discourages reporting of errors
Failure to report errors increases likelihood that error will be repeated
Also increases probability that errors may go undetected until consequences very severe

## Slide 35
Incident reporting systems:
internal or external systems

## Slide 36
Incident reporting systems
Different types of incident reporting systems:
Internal reporting systems:
Reporting inside organization (e.g. local incident reports)
External reporting systems:
Reporting outside organization (e.g. web-based systems)
Notes: This is another way of differentiating incident reporting systems. Again, examples of both will be given later in the module.

## Slide 37
Internal reporting systems
Reporting of incidents within organization
Specific in relation to intra-organization …
    … procedures
    … equipment
    … characteristics
“Lessons to learn” become more direct and explicit
Follows up management of actual patients affected by the incidents
Should evolve locally, but could be aided from the outside
Notes: Focussing now on the internal reporting systems (local incident reporting systems), the characteristics of these systems are given here. Local systems should evolve locally to best meet the needs of the organisation, but could evolve in a way that is open to comparison with best practice in local incident reporting elsewhere.

## Slide 38
Internal reporting systems
What type of information is it useful to include in an internal incident reporting system?
Administrative information (signatures, etc.)
Patient information (patient name and number, etc.)
Incident information (description of event, cause of event, when and where did event occur, who detected event (and when, where and how), estimation of deviation, clinical significance, contributing factors, etc.)
Action information (corrective action (and who is responsible for this), preventive action (to prevent recurrence), information to communicate (to patient, physician, authorities), etc.)
Notes: This slide contains some suggested elements of a local incident reporting system – based on a European survey (22 radiotherapy clinics) by the ROSIS group. It is not intended to be complete.

## Slide 39
External reporting systems
Reporting of incidents outside organization
“Lessons to learn” will come from a bigger pool of events
An incident in another hospital can lead to identification of the hazard before a similar incident is realised locally
With a more extensive pool of events, safety-critical steps in the radiotherapy process can be identified
A general culture of safety-awareness can be created by making this information available
Notes: Focussing finally on the external reporting systems, the characteristics of these systems are given here.

## Slide 40
IRID Ionising Radiations Incident Database
Supporting organizations:
National Radiological Protection Board (NRPB) – UK
Health and Safety Executive (HSE) – UK
Environment Agency – UK
Some characteristics:
National database (UK) of (non-nuclear sector) radiation incidents
Excludes patient exposure incidents
http://www.irid.org.uk/
Notes: Starting with this slide are a few examples of external incident reporting systems.

## Slide 41
ENR Event Notification Reports
Supporting organization:
Nuclear Regulatory Commission (USA) (regulates US civilian use of nuclear materials)
Some characteristics:
Mainly industrial events, but also some medical (cobalt and brachytherapy)
Feedback as preliminary notification reports
http://www.nrc.gov/

## Slide 42
RELIR Retours d’Expérience sur Les Incidents Radiologiques
Supporting organizations:
French Radiological Protection Society (SFRP)
… + Institute Curie (France)
Some characteristics:
As IRID, but also accidental medical exposure
Incidents assessed by radiation protection specialist for the specific activity: “Moderator”
http://relir.cepn.asso.fr/

## Slide 43
AHRQ WebM&M Morbidity & Mortality
Supporting organization:
Agency for Healthcare Research and Quality (AHRQ) US
(part of US DoH)
Some characteristics:
National forum (US) on patient safety and health care quality
Features expert analysis of medical errors reported anonymously
Features interactive learning modules on patient safety and forums for online discussion
http://www.webmm.ahrq.gov/

## Slide 44
The RADEV system
Notes: The next section contains a number of slides where the IAEA system “RADEV” is presented.

## Slide 45
RADEV
Radiation Events Database
Supporting organizations:
International Atomic Energy Agency (IAEA)
National regulatory authorities in signatory states

## Slide 46
RADEV
The purpose of RADEV is to help prevent accidents or mitigate their consequences, and to help Member States, IAEA and other organizations to identify priorities in their radiation safety programmes and to facilitate efficient allocation of resources

## Slide 47
A centralized RADEV database is being established at IAEA’s headquarters in Vienna to …
… provide a repository of information on accidents, near- misses and any other unusual events involving radiation sources not directly involved in the production of nuclear power or its fuel cycle
… categorize events in a standardised manner to facilitate the search for events fitting particular profiles, the identification of causes and the lessons to be learned
RADEV

## Slide 48
A centralized RADEV database is being established at IAEA’s headquarters in Vienna to …
… provide a means to analyze trends in radiation events
… provide summary descriptions of events that can be used directly as training material
RADEV

## Slide 49
The scope of RADEV:
Events with actual/potentially significant radiation protection consequences and from which lessons can be learned:
    accidents
    near misses
    any other unusual* events
    (*e.g.: malicious acts, deliberate acts)
RADEV

## Slide 50
The scope of RADEV:
Include:
  Worker / public exposure
  Loss of control of sources (lost, found, stolen, illegally transported or sold)
  Patient exposure significantly different than intended
Exclude:
  Nuclear power plants, fuel cycle and weapons
  Transport
  Illicit trafficking of nuclear materials
RADEV

## Slide 51
RADEV
Data collection and dissemination
RADEV
International
RADEV
National
RADEV
National
IAEA
Notes: Data is intended to be kept centrally at IAEA headquarters, and be fed in and fed back to national nodes.

## Slide 52
Dissemination of information
RADEV
International
IAEA
Internet
Professional
Journals
IAEA
RADEV
Annual
Report
RADEV
Notes: Various inteded methods for dissemination of RADEV data.

## Slide 53
Information included in this searchable database:
  Description of event (free-text)
  Type of practice involved in event (drop-down list)
  Radiation source (drop-down list)
  Type of event (drop-down list)
  Causes (drop-down list)
  Lessons to learn (free-text)
  Response brought on by event (tick-box)
  Corrective action (free-text)
RADEV
Notes: This is the profile of information contained in RADEV.

## Slide 54
Notes: A screenshot of RADEV.

## Slide 55
The ROSIS system
Notes: The next section contains a number of slides where the “ROSIS” system is presented.

## Slide 56
http://www.rosis.info
ROSIS Radiation Oncology Safety Information System
Supporting organization during the development phase:
European Society for Therapeutic Radiology and Oncology (ESTRO)

## Slide 57
Proposed in 2001 by Mary Coffey and Ola Holmberg
Steering Committee:
Ola Holmberg, Project Leader, Malmö
Mary Coffey, Dublin
Joanne Cunningham, Research Assistant, Dublin
Tommy Knöös, Lund
Ingrid Kristensen, Lund
Torsten Landberg, Copenhagen
Håkan Nyström, Copenhagen
Ann Barrett, Norwich
Hans Svensson, Stockholm
Guy Francois, ESTRO, Brussels
ROSIS
Working
group

## Slide 58
Main objectives:
To establish a web-based system whereby radiotherapy incidents can be analysed in a systematic and objective way, and the information shared through web-access to a centralised database
To enable radiotherapy clinics to address safety issues before an accidental exposure occurs and to create a general culture of safety awareness by making information available on details of incidents, near-incidents and corrective actions, submitted on-line by other radiotherapy clinics
ROSIS

## Slide 59
Main objectives:
To define a hazard classification system and perform frequency analysis, leading to the identification of safety-critical steps in the radiotherapy treatment process where errors are likely to occur or be detected
To allow current best practice in incident reporting within medical as well as non-medical settings to be utilized in radiotherapy by identifying high-reliability organizations outside radiotherapy and the methods used within these organizations for incident and near-incident reporting, evaluation and feedback
ROSIS

## Slide 60
How does the reporting work?
  On ROSIS web-site: www.rosis.info
  First time: fill in “Department-form” (profile of dept.)
  Submit “Department-form” on-line
  A “Clinic-ID” (unique number) will be sent back
  When the clinic wants to submit an event / incident to ROSIS, the “Incident-form” is filled out and submitted on-line
ROSIS

## Slide 61
Notes: This is a composite screenshot of the ROSIS incident report form – filled in on line by registered clinics.

## Slide 62
Notes: Screenshot of the ROSIS interface on the web.

## Slide 63
How does the sharing of data work?
  On ROSIS web-site: www.rosis.info
  Click on “ROSIS Data” and “Goto queries”
  You can now do free searches in the entire database of reported incidents (>1000 in 2007) or make specific queries such as “Where in the process / By whom / Through what type of QA procedures – are incidents discovered in radiotherapy”. Anyone can search, but the clinic needs to be registered to submit
ROSIS

## Slide 64
ROSIS
Incident: 265
Short description: A wedge was not used where it should have been
Treatment modality: External
Discovered by: Therapist (trt unit)
How Discovered: Found at time of patient treatment
Anyone affected: One patient
Severity: Light
Description: Patient received 2 fractions of treatment without a wedge in the lateral portions. The treatment had been planned with wedged laterals.
Cause: The accessory had not been entered into the verification system and checking procedures failed to pick up the error.
Suggestion: Reiterate the room checking procedure and the control monitor checking procedure.
Notes: This slide contains an example of an incident that has been submitted to ROSIS – just to give a flavour of what is contained within the database.

## Slide 65
ROSIS – first 200 reports
Who discovered the reported event?
Mainly Radiation Therapists on the treatment units
Notes: The following few slides contain some summarised information of the profiles of the first 200 reported incidents to ROSIS – their characteristics in relation to certain questions.

## Slide 66
ROSIS – first 200 reports
At what “check-station” did the discovery happen?
Mainly at chart check. However, most discoveries through “vigilance” at time of treatment.

## Slide 67
ROSIS – first 200 reports
What is the estimated outcome or potential outcome for the patient / person(s) affected? (Response)
… out of 128 incidents (and no near-misses)

## Slide 68
ROSIS
A hazard identification system is being worked on, based on the submitted reports and 910 reports on paper (retrospective) previously submitted by clinics Categories have subcategories, etc. Frequency analysis will be attempted
Notes: Work in progress (2004): A hazard identification system with corresponding frequency analysis for radiotherapy.

## Slide 69
Summary
Incident reporting systems represent one of the layers in the multilayered prevention of accidental exposures
Other layers, such as procedures, etc, are also needed
There are mandatory / voluntary and internal / external incident reporting systems
Different types of systems have different foci, but are all important

## Slide 70
Summary
Departments should work to develop a safety culture that encourages the reporting and addressing of errors, incidents, and potentially dangerous circumstances
Incidents should be investigated promptly and action taken to:
  Mitigate the effects of the incident
  Educate the staff on the causes and effects of the incident
  Establish procedures to prevent the occurrence of a similar incident

## Slide 71
References
Department of Health (UK) – An organization with a memory (2000)
Croft J R, Crouail P, Wheatley J S. Review of the development of incidents databases and feedback mechanism: IRID, RELIR, EURAIDE and RADEV
Holmberg O, Coffey M, Knöös T, Cunningham J and the ROSIS Steering Committee. ROSIS - at the conclusion of ESQUIRE. ESTRO News 57: 15-16 (2004)
Holmberg O. Ensuring the intended volume is given the intended absorbed dose in radiotherapy - Managing geometric variations and treatment hazards (ISBN 91-628-6002-X) (2004)